目 录

一、脑机接口硬件	1

（一）侵入式脑机接口硬件	1

（二）非侵入式脑机接口	11

二、脑机接口软件	17

（一）侵入式脑机接口软件	17

（二）非侵入式脑机接口软件	25

三、脑机接口外设	29

四、脑机接口临床应用	29

（一）侵入式脑机接口应用	29

（二）非侵入式脑机接口临床应用	56

参考文献	77

脑机接口

脑机接口（Brain-Computer Interface，BCI）旨在不依赖于外周神经与肌肉，在大脑与外部设备之间建立一种直接的、全新的信息沟通和控制通道，从而实现脑与机器之间信号的双向交互。脑机接口作为塑造未来产业的颠覆性技术，具有高科技、高效能、高质量特征，已成为全球科技前沿热点，是引领未来科技发展与产业竞争的战略制高点。脑机接口衍生出大脑解析、控脑、闭环神经调控等一系列变革性技术，可能深刻改变人机交互与认知增强的模式，对经济社会发展产生重大影响，有望孕育出未来产业的新增长点。

# 一、脑机接口硬件

# （一）侵入式脑机接口硬件

侵入式脑机接口系统的核心硬件，特别是直接与神经组织形成界面的电极和专用芯片，构成了整个技术体系的物质基础。当前全球技术竞争格局呈现出明显的梯次分布特征，欧美等国通过数十年持续投入已构建完整的技术生态，国内通过在高通量电极、生物融合电极、快速产业化等方向实现突围。

- 电极

侵入式神经电极根据植入位置不同，主要可分为脑皮层电极（Electrocorticographic Electrode, ECoG）、血管介入式电极（Endovascular Electrode）和皮层内电极（Intracortical Electrode）[1]三种技术路线。其中脑皮层电极和血管介入式电极不直接进入脑组织内部，手术损伤较小，但时空分辨率低，传输信息量有限；皮层内电极直接植入到脑组织内部，手术风险更高，但能够精准记录单个神经元级别的神经信号，直接反应大脑生理状态和信息编码，能真正实现与大脑的高时空分辨率、高带宽信息交互。

1.1 脑皮层电极

脑皮层电极主要分为硬膜外电极和硬膜下电极。

硬膜外电极。硬膜外电极放置于大脑硬脑膜外，由于避免穿透硬脑膜，侵入性较小，手术安全性较高。但由于与脑组织距离较远（2-5mm），只能记录大脑中大量神经元群体活动的平均信号（即场电位信号）。同时，由于神经信号在穿过硬膜过程中发生进一步衰减，导致硬膜外电极时空分辨率较低，主要记录0.5-100Hz低频场电位信号。国内清华大学洪波团队和博睿康公司联合研发的无线硬膜外脑机接口系统NEO，通过将8通道电极放置于硬膜外，成功实现颈髓损伤患者抓握功能的准确分类。

硬膜下电极。硬膜下电极通过硬膜开口置于硬膜下，直接贴附于大脑的表面，手术创伤相对较大[2]，但具有更高时空分辨率，可捕获80-500Hz信号。美国Precision Neuroscience开发了Layer 7硬脑膜下柔性电极，通过硬膜切口将电极放置于大脑表面，在1例患者脑内植入了4096通道的柔性硬膜下电极产品，并完成了术中信号监测。国内脑虎医疗联合华山医院为1例语言区占位肿瘤癫痫患者植入了自研的256通道硬膜下柔性电极，患者在术后5天实现了142个常用汉语音节下71%的解码准确率。

1.2 血管介入式电极

血管介入式电极通过血管介入手术将电极部署于脑血管腔内，无需开颅，手术风险相对较低，但由于血管的屏蔽作用，时间和空间分辨率均低于脑皮层电极。另外，血管介入式电极的植入通道数低，且植入位置极大地受到血管走向限制，长期植入还可能导致血栓风险。澳大利亚Synchron公司的Stentrode产品采用血管内微创植入技术，将16通道支架式电极通过颈静脉部署到患者的上矢状窦，已完成多例临床手术，帮助患者实现了移动光标、打字、控制机械臂等运动功能。美国哈佛大学Lieber团队研发超小型、超柔性神经电极[3]，植入大鼠内径小于100μm的动脉血管中，通过调整柔性电极的机械性能实现递送时对血管分支的选择性，实现嗅球神经元信号记录。国内中科院脑科学与智能技术卓越创新中心赵郑拓团队通过颈部静脉入路，在绵羊动物模型上将柔性电极和穿刺装置递送到达窦汇处，并穿过血管壁植入视觉皮层[4]，实现大动物模型上免开颅手术验证和神经元信号记录。

1.3 皮层内电极

皮层内电极直接植入脑组织内部，能够记录单个神经元级别的高时空分辨率信号，是实现高带宽脑机接口交互的核心技术路径，但手术风险更高且长期稳定性易受胶质增生影响。主要有柔性线性电极、柔性集束电极和柔性网状电极等类型[5]。

柔性线性电极。柔性线性电极由微米级厚度柔性基底与百纳米级厚度的金属导电层构成，通过钨丝引导针或者可降解包裹材料辅助进行植入。柔性线性电极具有和脑组织更加匹配的力学性能，电极通道数高且植入损伤小，是侵入式柔性神经电极的主流电极形态。美国德克萨斯大学奥斯汀分校Chong Xie团队通过减小神经电极尺寸、选用柔性高分子材料等，制成了超小超柔性线性神经电极，可在小鼠脑内稳定工作，达到长期记录无瘢痕的目的[6-7]。临港实验室李澄宇团队联合北京脑科学与类脑研究所方英团队[8]研发高密度1024通道柔性神经流苏电极，其截面尺寸达到了神经元轴突的级别，是目前全球通道数最高且尺寸最小的柔性线性电极。北京大学段小洁团队[9]开发“神经卷轴”探针，其长度可在1-10cm范围内灵活设计，从而能够适应从啮齿类到灵长类动物的记录需求。浙江大学宋吉舟团队利用形状记忆聚合物材料与可折叠分支结构结合的设计，研发的柔性电极能够以折叠状态下植入，进入小鼠脑内后能够自动软化并展开[10]。美国Neuralink公司1024通道无线脑机接口产品获得了FDA试验用器械豁免（Investigational Device Exemption, IDE）[11-12]，通过机器人辅助植入技术，将64根柔性线性电极精准植入患者大脑皮层内部，成功帮助患者实现了对外部设备的高精度意念操控，目前已完成超20例临床试验。上海阶梯医疗无线高通量脑机接口产品，采用无线供能方式最高支持64通道信号采集，是国内首个完成型检的皮层内脑机接口产品。北京智冉医疗无线高通量脑机接口产品最高支持256通道，配合自研深度学习解码算法，能够实现高精度脑控光标。该公司研发的脑深部高通量柔性电极，同时支持采集与刺激功能，已与浙大二院合作完成了2例患者的脑深部电极植入手术，首次实现了对人脑深部核团神经元信号的高通量信号采集。

集束柔性电极。集束柔性电极通过热拉伸和微纳加工技术，将微丝、光纤、微管等材料集成为多功能侵入式器件，具备多模态优势，但通道数相对较低，尺寸较大，对脑损伤大。加州大学Sirbuly团队[13]研发的多模态同轴集束电极，将光刺激和电生理记录功能结合，通过在微光纤表面沉积氧化铱和聚（3,4-乙烯二氧噻吩）等材料，降低对神经组织的损伤和免疫反应。麻省理工学院Xuanhe Zhao团队[14]研发了集光刺激、电生理记录和药物输注于一体的集束电极，通过结合柔性的水凝胶基质，突破电极与神经组织力学性能不匹配的问题。

网状柔性电极。网状柔性电极采用多孔网状结构设计，通过精确调控网孔尺寸和整体结构，能灵活调整其机械强度和动力学特性，以适应不同神经组织环境，通过引导针辅助、自组装材料包裹或直接注射等方式植入目标区域。独特的网状结构能够与神经元形成空间上的无缝交织网络，从而有效降低免疫排斥反应和神经信号漂移，同时具备随大脑发育动态调整能力。目前网状柔性电极仍面临植入损伤较大和电极密度较低的局限性。哈佛大学Lieber团队[15]开发的拟神经元电极（NeuE），采用立体定位注射技术，将拟神经元电极精确地植入小鼠大脑的目标脑区，显著减少了免疫反应和组织损伤。哈佛大学Jia Liu团队[16]研发的网状电极同样具有超柔性和高生物相容性，电极尺寸和形状为匹配生物组织的微观结构而设计，在组织发育过程种将纳米电子电极直接整合进生物组织，利用生物组织的自然发育机制，使得电极与神经组织间从初期的嵌入到界面逐渐优化，从而形成稳定的电生理连接。

- 芯片

芯片是侵入式系统的核心部件，承担神经信号的高精度采集、实时处理、智能解码及闭环调控等关键功能。当前国际技术发展呈现“微型化、高集成、低功耗”三大趋势，国内在专用芯片研发方面也取得突破性进展，但整体产业化水平仍在发展。

国外侵入式脑机接口重点聚焦无线全侵入式、高信噪比、高通量、高时空精度、高带宽全频谱脑电信号检测和闭环刺激调控，因而对芯片的功耗、通道数、噪声、共模抑制比、电源抑制比、采样率、信号带宽无线传输能力和刺激调控功能都有较高要求。基于30-40年研发历史积累，采用从学术界到商业界逐次推进的方式，实现从科研到产业的转化。例如早期犹他大学团队[17]、密歇根大学团队[18]、布朗大学团队[19]，到后来斯坦福大学团队[20]、华盛顿大学团队[21]、欧洲微电子研究中心（IMEC）[22]，到现在Intan公司、NeuroNexus公司等。

美国Intan Technologies公司是当前世界市场占有率最高的侵入式脑机接口信号采集芯片公司，其产品可以实现神经脉冲信号、场电位和脑皮层电图等神经活动信号的采集，可对神经元施加可控参数的电脉冲刺激。其代表性RHD2164信号采集芯片，内部集成放大器、时分复用器、模数转换器，采用电容反馈的低噪声跨导放大器（OTA）技术，可实现64同步信号采集通道，2.4µVrms输入噪声，30kSps采样率，0.1Hz到20kHz可调带宽，16比特信号精度，3.3V通用CMOS电压供电，芯片功耗偏大，总功耗约为53.1mW（30kSps/ch采样率、10kHz带宽、3.3V供电情况下），即830μW/通道[23]。RHS2116芯片此基础上添加了刺激功能，可实现神经信号的采集和刺激，但其通道数仅为16[24]。

欧洲微电子研究中心自2015年开始将预放大器直接集成在神经电极上，设计出一系列超高通道数（约1000通道）、高集成度的神经记录芯片，并完成商业化成立Neuropixels公司。Neuropixels公司Neuropixels 1.0芯片达到能同时从384通道采集信号采集，并通过时分复用实现从966个电极采集信号。Neuropixels 2.0系统同步采集通道仍为384个，但时分复用的电极位点拓展到了1280/5120。该芯片同样使用电容反馈的低噪声跨导放大器技术，通过和130nm CMOS工艺和后CMOS处理技术实现电极和芯片一体化集成，形成了单一一体化芯片电极器件，实现384同步通道数、49μW/通道功耗、6.7μVrms噪声、60dB共模抑制比、70dB电源抑制比、30kSps/ch采样率、0.5Hz到10kHz可调带宽、10比特信号精度、1.8 V通电压供电[25]。

Neuralink公司2020年研制出一款256通道超小尺寸芯片，采用模块化设计，由4个64通道模块组成，技术指标达到了256同步通道数，纯芯片总功耗6mW（23.4μW/通道）、5.9μVrms噪声、19.3kSps/ch采样率，3Hz到27kHz带宽，10比特信号精度。通过多块芯片集成，形成两种不同配置的系统：一个包含1536通道（System A, 550mW），一个包含3072通道（System B, 750mW）的微型脑机接口器件，支持实时无线传输及供能[26]。2021年通过在猴子脑内植入芯片成功让其实现了对运动信息的解码，2023年获得美国FDA IDE认证开始招募临床受试者，并在次年完成3例植入实验，均取得了较好的数据采集结果和实时控制的脑机接口实验。2025年6月，Neuralink公司发布了一种1680通道的神经调控芯片，具有1680个刺激通道及16个采集通道，集成脉冲探测算法功能以及刺激调控功能，可用于高精度视觉皮层神经调控，目标是帮助患者恢复初步视觉功能。

美国老牌侵入式脑机接口器件和应用公司Blackrock Neurotech已实现512通道商业化应用，其核心芯片是一款自研的128通道神经脉冲信号采集芯片[27]。美国NeuroNexus公司推出了芯片电极一体集成的有源电极阵列SiNAPS Pixel Probes，集成硅基CMOS探针及全套一站式记录系统。其信号放大、低通滤波和多路复用读出等有源电路直接集成在每个电极像素下方，提供超高密度神经信号采集，有256通道和1024通道两个版本，每个通道6μW，20kHz采样率，6.5μVrms输入噪声，电极点大小14μm×14um，pitch 29μm，电极长5.6mm，宽88μm，厚50μm，间隔560μm[28]。

国内在侵入式采集芯片和刺激芯片研究还处于前期投入和成熟产品迭代验证阶段，已有产品上市或正在进行相芯片研发投入的企业包括神芯科技、矩智科技、全澜科技、微灵医疗、芯智达、阶梯医疗、脑虎科技、华为等。神芯科技在2025年推出128通道采集芯片（已上市），具有刺激功能，主要对标Intan公司产品，其在功耗和体积上有明显的优势，但噪声为3.5μVrms略高于Intan芯片2.4μVrms[29]。上海矩智科技研发神经信号采集芯片MTXNEP256，可进行256通道的脉冲信号采集，其体积和功耗也优于Intan，其噪声较高约7μVrms，精度略低约为12bit，目前处于对外样片试用阶段。同时，推出了320通道高通量刺激芯片以及多款低通道生物信号检测芯片[30]。宁矩科技在2020年左右计划研制推出128通道采集芯片。临港国家实验室也在以最终产品落地为目标研发高通量神经信号采集芯片。

在半侵入式微创脑机接口应用方面，深圳微灵医疗有限公司正研发一款高通量低功耗的硬膜外神经信号采集芯片，目前其第一代芯片已经完成流片和测试[31]。复旦大学于2020年研制的64通道无线脑电采集芯片，兼顾无线信号传输和无线能量传输模块，系统功耗约2.5mW，但通道数较少，且其噪声比同类国外商业产品的噪声要高2-3倍，前端放大器输入动态范围较小，容易造成放大器的饱和失真，且其采样精度仅为10bit，比同类产品要低2-6个比特。浙江大学提出一款无源128通道无线记录芯片，采用近距离背板传输无线供电，通信端提出新颖的多载波正交调制技术实现多通道数据同时传输，通信速率达20.16Mb/s，能效低至0.8pJ/bit，整个芯片记录功耗仅为9.87μW/ch。然而，其记录噪声偏高，有效精度较低，动态范围仅为52dB，无法有效抑制伪影。

- 集成系统

侵入式脑机接口系统技术已形成"科研仪器-植入设备-临床治疗"的完整产业链，当前技术发展呈现微型化、无线化、智能化三大趋势，推动脑机接口从实验室走向临床应用。

欧美国家在系统高度集成、新型脑机接口、临床应用等方面有重点突破。美国Neuralink研发了尺寸仅为23mm×8mm的1024通道全侵入式微型系统，目前已在全球范围内完成超20例的临床试验，患者可成功通过意念控制电脑、玩游戏，甚至操控机械臂完成更复杂的动作。美国Paradromics研制了可记录1600多个独立神经元信号的无线高速率脑机接口产品Connexus直接数据接口（Direct Data Interface，DDI），在2025年12月获得FDA研究性器械豁免（IDE）批准，已于2026年1月正式启动Connect-One研究，计划招募2名严重运动障碍患者，在密歇根大学等三家中心评估言语功能恢复能力。美国医疗科技公司Medtronic研发了全球首款具有感知技术的商用脑深部刺激器——智能可感知脑起搏器Percept™PC，已广泛应用于全球范围内帕金森患者临床治疗。美国微创生物电子植入物研发商Motif Neurotech开发了仅穿透头骨即可实现大脑刺激调控的数字可编程脑外治疗仪（Digitally Programmable Over-brain Therapeutic，DOT）[32]。美国精密神经科学公司Precision Neuroscience构建了一种模块化、可拓展的脑机接口平台——第7层皮层接口，高通量薄膜电极阵列和微创手术植入系统使得无需进行开颅手术即可实现皮层表面电极植入和双向通信，目前已在50多名患者中完成植入，帮助瘫痪患者完成打字、玩游戏等任务。澳大利亚Synchron、美国医疗技术公司Forest Neurotech等研发了介入式血管内运动神经假体系统、微创超声等新型侵入式脑机接口系统。

国内相关企业已完成了不同类型的微系统集成，目前正陆续开展脑疾病诊疗、运动康复等临床研究。清华大学与博瑞康股份有限公司研制了电极约为两枚一元硬币大小的无线微创植入脑机接口系统NEO[33]，已完成共计36例临床植入试验，并获得国家药监局批准注册上市。北京品驰医疗设备股份有限公司开发了尺寸为12mm×70mm×61mm的脑深部刺激系统，可应用于帕金森病、肌张力障碍等疾病治疗，目前全国累计植入已超过2.5w次。景昱医疗科技（苏州）股份有限公司开发了规格为50mm×48mm×10mm的可充电脑深部电刺激系统，已获批上市正式投入临床应用。阶梯医疗自主研发的256通道无线高通量侵入式脑机接口系统（WRS02）已完成3例临床植入，并成功验证脑控交互功能，已进入国家药监局创新医疗器械"绿色通道"的侵入式脑机接口产品，计划2026年中启动大规模多中心注册临床试验，年内完成约40例患者入组。北京脑科学与类脑研究所牵头研制了半侵入式、无线全植入的智能脑机接口系统（北脑1号），系统搭载‌128通道柔性高密度电极‌，在已完成的临床植入案例中，设备‌有效通道数保持率超过98%，信号质量可长期稳定，可实现‌单字解码时延低于100毫秒，‌渐冻症患者可通过该系统实现62个常用字词的‌解码输出‌。脑虎科技全自主研发的256 通道高通量侵入式柔性脑机接口，先后开展高精度实时运动解码和语言解码临床试验研究，成功实现了“脑控”智能设备和“意念对话”。杭州佳量医疗科技有限公司研制的面向难治性癫痫的侵入式自响应神经刺激系统EpilcureTM已开启多中心、对照临床试验，完成超百例注册临床试验，正在准备提交注册申请。

# （二）非侵入式脑机接口

非侵入脑机接口电极和芯片正经历从单一信号采集向“材料-硬件-算法-应用”全链条创新的深度变革。

- 电极

电极与头皮直接接触，决定了采集脑电信号的质量和空间分辨率，目前商用电极仍以具有较低阻抗的湿电极方案为主（Neuroscan、g.tec和Ceribell等公司）。宾夕法尼亚州立大学开发的毛发状水凝胶半干电极在24小时的连续佩戴中信号质量保持稳定，中科院半导体所开发的预凝胶化半干电极提升了穿戴效率。西北大学Rogers团队开创了表皮电子的先河，其开发的基于聚酰亚胺和Ecoflex的大面积柔性脑电电极可与头皮保形接触，在MRI环境下，该电极能够同步进行高质量EEG信号采集。微针阵列干电极在13届MEMS会议上首次提出，采用硅基MEMS工艺加工而成，近年来随着新型材料和加工工艺的开发，佐治亚理工大学和北京大学所开发的微针阵列电极表现出良好的可穿戴特性和低阻抗优势。上述半干电极和干电极相比于湿电极具有长时程稳定采集的优势，且通过材料和加工工艺创新，适配弯曲皮肤表面，优化界面特性，具有更优良的可穿戴特性，有望解决非侵入式脑机接口长期高质量脑电信号采集的挑战。目前非侵入式脑机电极空间分辨率尚未达到理论和仿真研究的极限（电极中心距3.3mm），进一步提升脑电信号采集的空间分辨率，可实现脑电信息的充分采集，助力脑机接口技术快速发展。

1.1 湿电极

澳大利亚Compumedics Neuroscan公司研发的基于湿电极的脑电帽，通过导电膏形成电极与皮肤间的导电通路，显著降低电极-皮肤接触阻抗（＜10 kΩ），是当前脑机接口研究主流脑电采集方案，广泛应用于P300、SSVEP、运动想象等脑机范式[34]。

奥地利g.tec公司开发了基于湿电极的高密度“穿山甲”电极，其电极中心距为8.6mm，电极直径为5.9mm，覆盖全头可实现1024通道脑电信号采集，显著提高了EEG采集的空间分辨率，是目前商用的最高密度电极，通过监测大脑躯体感觉/运动诱发电位，实现无创精准定位大脑皮层中央沟[35]。

美国Ceribell公司推出的床旁EEG系统，基于湿电极实现脑电信号采集，借助人工智能技术，能够对癫痫活动实时评估，有望重塑急诊环境中的癫痫监测技术，Ceribell系统已经获得了美国 FDA 510（k）许可，2024年10月在美国纳斯达克上市[36]。

1.2 半干电极

柔灵科技开发的水凝胶脑电贴片，可用于便携式睡眠监测，其监测准确率区间约为脑电帽监测的86~94%之间[37]。

中科院半导体所研究团队提出了一种预凝胶化电极，将水凝胶预先放置在Ag/AgCl电极上，缩短了电极安装时间，阻抗约50kΩ–100kΩ，实现了与湿电极相当的信号采集性能。在40目标SSVEP任务中，该电极的分类准确率与信息传输率和湿电极无显著差异[38]。

美国宾夕法尼亚州立大学研发了一种新型毛发状电极，通过3D打印技术制备水凝胶电极，并利用生物粘附墨水直接粘附在头皮上。10Hz下阻抗为16.72kΩ，并且在超过24小时的连续佩戴中信号质量无任何下降[39]。

1.3 干电极

美国西北大学Rogers团队开发了一款大面积柔性表皮电子接口，该电极采用岛-桥结构设计，使用聚酰亚胺和Ecoflex作为基底材料，通过去除铁磁性材料实现了MRI兼容。30 Hz下阻抗约30kΩ。在MRI环境下，该电极能够同步进行高质量EEG信号采集，用于监测注意力、决策等认知活动[40]。

中国科学院深圳先进技术研究院提出了一种基于可拉伸粘弹性材料的干电极，电极由经过导电材料与粘弹性材料改性的聚硼硅氧烷基质组成，100Hz下归一化阻抗约为132kΩ·cm2该基质带来的半流动特性使电极能穿透头发，与皮肤保形贴合。在EEG信号监测中，该电极能稳定监测α波，并能清晰地检测到听觉诱发电位[41]。

中国科学院半导体研究所开发了具有弹性金柱结构的16通道刚性超高密度脑电电极阵列，电极中心距仅5mm，1Hz下阻抗约100kΩ，并验证了阵列通道之间的信号具备良好的独立性，但其刚性衬底设计导致佩戴舒适性较差，难以扩展至全头[42]。德雷塞尔大学和宾夕法尼亚大学联合开发了基于MXene的阵列式柔性干电极阵列，电极中心距为6mm，10Hz下阻抗约56kΩ，相比于单个Ag/AgCl电极，该电极阵列能更加充分地采集Beta波段空间信息[43]。

美国约翰・霍普金斯大学开发的碳纳米管复合电极，通过化学气相沉积技术在柔性基底上生长垂直排列的碳纳米管阵列，电极阻抗较传统银/氯化银电极降低60%，且耐汗液侵蚀能力提升3倍。该电极在长期监测中显示，α波信号幅度波动＜8%，已进入临床前试验阶段，用于监测阿尔茨海默病患者的睡眠纺锤波异常[44]。

美国佐治亚理工学院研究团队开发了一种基于聚酰亚胺的硬质微针电极，借助蛇形线结构实现岛-桥可拉伸结构，提升可穿戴性。其电极阻抗约30kΩ，EEG信号采集信噪比仅为13.7±2.6dB。在运动想象的分类任务中，其分类准确率为93.22±3.33%，通过集成VR设备，该系统成功应用于实时无线控制VR游戏[45]。此外，该团队提出了一种可放置于毛囊间的微针电极，电极通过紫外线复制模塑工艺制备，并选用PEDOT:Tos导电聚合物，该电极在站立、行走、跑步状态下进行了基于SSVEP的任务，信号分类准确率高达96.4%[46]。

北京大学研究团队提出了基于聚酰亚胺的柔性微针阵列电极，该电极通过两步微模塑工艺制造，并利用导电聚合物PEDOT:PSS对表面进行修饰，10Hz下阻抗约4kΩ，归一化阻抗约为临床湿电极的1/250，并首次将微针电极用于临床研究中，开展了健康被试和OSAHS患者的睡眠临床研究[47]。团队还提出了一种可拉伸微针阵列电极，该电极由刚性环氧树脂微针和可拉伸的SEBS基底组成，组装在可拉伸的聚氨酯贴片上，进一步提升可穿戴舒适性。电极10Hz下阻抗为1.70±0.18kΩ，1kHz下阻抗为1.09±0.09kΩ，经过300次拉伸疲劳测试后阻抗变化低于10%，在自由运动条件下能连续8小时进行高质量EEG信号采集[48]。

- 芯片

高精度脑电采集芯片仍是非侵入脑机接口的主要需求，我国经过十四五期间的部署和发展，国产高精度采集芯片已经获得部分成效，基本实现8通道24位脑电采集芯片的研发；柔性电子技术突破传统刚性限制，柔灵科技的12μm厚度有机电化学晶体管芯片、韩国3mm³无线胶囊芯片等实现动态场景下28dB信噪比的稳定信号采集；多模态融合芯片通过MR增强交互、5G云端协同（华为Balong BCI-600）等架构，将脑控机械臂延迟压缩至8ms，定位精度达0.5mm。材料科学的突破尤为显著，清华大学忆阻器阵列、东京工业大学生物相容性水凝胶电极等新材料，推动芯片能效比提升100倍以上，且解决了长期监测中的生物相容性难题。美国Neuralink的Blindsight芯片获FDA批准，支持非侵入式脑控机械臂；德国西门子医疗128通道工业级芯片实现＜10ms延迟的实时监测；日本Looxid Labs的VR脑电传感器接入Steam平台，推动消费级市场破冰。然而，头皮衰减导致的信号质量瓶颈（有效信号＜100μV）、多模态融合的时空异步问题，以及跨学科协同的长周期挑战，仍需在材料改性、架构创新与生态整合中持续突破。

2.1 单一高精度脑电采集芯片

脑电信号感知与采集关键技术方面，作为脑机接口领域的关键部分，近年来在全球范围内取得了显著进展。目前脑机接口芯片仍以国外芯片为主，特别是非侵入式的脑电采集芯片主要是使用美国TI公司的ADS1299芯片，国内部分企业进行相关研究并取得阶段性成果。2024 年，中电云脑公司联合天津大学、上海交通大学等单位基于国产工艺和正向设计，已成功完成八通道国产脑电采集芯片开发和设计，该芯片拥有完全自主知识产权，能够实现对微弱脑电信号的高精度采集，在模拟数字转换器动态范围、放大器低噪声技术和低功耗设计技术上均有所提升，具备与市面上现有商用八通道采集芯片的兼容性，并支持菊花链扩展和双电源配置，优化了用户体验，该芯片已完成整体性能测试，部分关键指标超越ADS1299。

2.2 神经信号计算芯片

中电云脑公司联合脑机海河实验室团队开展了可编程电生理信号智能编解码计算芯片研究，该芯片采用架构安全的RISC-V处理器核心，资源和高速接口方面，等效逻辑单元容量为94K， BRAM容量从4.9Mbit提升至5.73Mbit，DSP数量从220个提升至240个；内嵌了特征运算加速模块，提供了高效的神经电生理信号解析能力，解决了神经电生理信号干扰大，运算量大且复杂的计算问题。

清华大学集成电路学院与天津大学脑机海河实验室合作，提出了基于忆阻器类脑计算芯片的新型脑机接口解决方案。团队研制了一款128Kb规模的忆阻器芯片作为自适应脑电解码器，构建了完整的脑机接口系统（包含脑电采集、解码和执行器），提出了单步解码策略和交互式更新框架，实现了长期稳定的高效高精度脑电解码与脑机协同演进[49]。

加州理工学院研发的ASIC芯片采用近存计算架构，将模数转换器（ADC）与神经解码单元集成在同一硅片，数据搬运能耗降低80%[50]。

2.3 神经运动接口芯片

Meta的Brain2Qwerty深度学习模型通过非侵入式脑机接口解码句子生成，准确率高达80%。2025年Meta宣布开发出一种基于表面肌电图（sEMG）的通用非侵入性神经运动接口发表于《自然》杂志的论文中，该接口以腕带形式呈现，名为 sEMG-RD，采用干电极设计，包含 48 个电极，配置为 16 个双极通道，有四种尺寸以适应不同腕围，可精准捕捉手腕、手部和前臂肌肉的电信号，甚至能检测到单个运动单位动作电位，让“意念控制”成为现实[51]。

- 集成系统

美国纽约州卫生署Wadsworth研究中心牵头开发了已被广泛用于世界范围内约500个实验室的BCI2000通用脑—机接口软件平台，其灵活的机制可支持各类脑机接口研究领域应用，能匹配不同采集设备、设计算法及实验设计范式，快速建立实时脑机接口系统。脑机交互与人机共融海河实验室发布脑接机口领域综合性开源软件平台MetaBCI，针对数据分布散乱、算法复现困难、在线系统效率低等共性问题，开发了通用的算法框架，利用了双进程和双线程提高在线实时效率，有助于降低构建脑-机接口系统的技术门槛，减少研发成本，促进转化落地，为脑机接口全链条系统开发及应用提供了统一框架[52]。

# 二、脑机接口软件

脑机接口编解码是了解脑、控制脑的关键技术，通过分析和处理大脑产生的生物信号（如脑电波EEG），将其转换为可由计算机或其他外部设备识别和执行的指令或命令，从而实现用户通过思维与设备交互的目的。

# （一）侵入式脑机接口软件

解码技术作为脑机接口核心竞争力，直接影响系统的控制精度、响应速度和适用场景。针对侵入式脑机接口脑信号的高噪声、非稳态、强个体差异等显著挑战，分别发展了基于强化学习的解码、基于不同空间和时间尺度的解码以及跨脑区预测神经信号的解码。

- 基于强化学习的解码

强化学习（Reinforcement Learning, RL）在脑机接口领域的应用取得了长足进步，通过优化RL算法不断提升BCI系统的解码性能和适应能力。研究主要聚焦于基于Q-value的算法和Actor-Critic算法。随着任务复杂度和用户群体的增加，个性化学习与迁移学习逐渐成为热点方向。

1.1 基于Q-value的算法

Q-learning是最早应用于BCI的单步解码算法之一。2009年，DiGiovanna等人[53]首次将其用于解码大鼠脑信号，成功实现了机械臂控制，验证了其潜力。为了提升Q-learning对神经信号的处理能力，2015年，Wang等人[54]提出了注意力门控强化学习（Attention-Gated Reinforcement Learning, AGREL），该方法通过选择性关注关键神经特征优化动作选择，在猴子操控光标的任务中显著提高了成功率，凸显了注意力机制的价值。为了解决AGREL算法会陷入局部最优问题，Wang等人[55]在2017年提出了一种基于核强化学习的解码器（QAGKRL），通过在再生核希尔伯特空间（RHKS）进行神经信号解码以达到全局最优。2019年，Zhang等人[56]通过在先RKHS中聚类神经模式、再进行动作解码的方法，加速了算法的学习过程并提高了解码精度。针对传统RL输出离散动作的不足，为了更加合适地去解码神经假肢的连续运动轨迹，2023年，Zhang等人[57]将核强化学习与卡尔曼滤波相结合，开发了一种处理连续神经动作信号的解码器，通过平滑动作状态预测提升了实时控制任务的稳定性，为BCI的实用化提供了新思路。

时序差分学习（Temporal Difference learning, TD）是一种经典的多步强化学习方法，通过预测未来奖励的期望值和每一步动作之间的关系来优化策略。传统TD-learning采用线性逼近，为了更加适应BCI里的非线性神经信号，Bae等人[58-59]提出了Kernel TD方法。通过核函数来逼近值函数，KTD显著提升了准确率，并在猴子多步骤到达任务中获得成功应用。为了解决环境噪音问题，Bae等人[60]随后引入相关熵Correntropy作为一个新的目标函数，从而使得KTD算法在BCI应用中的鲁棒性更强。为进一步提升多步骤任务中的解码性能，Shen等人[61]在2022年提出把QAGKRL和KTD方法结合的新框架，该方法通过核函数逼近值函数，利用时序差分学习优化策略，同时结合内侧前额叶皮层mPFC的中间反馈信息，显著提高了复杂任务的准确率和稳定性。

1.2 Actor-Critic算法

Actor-Critic算法结合Actor（动作选择）和Critic（策略评估），在BCI中实现了更稳定和灵活的学习过程，根据Critic来源差异可以分类为算法生成型和被试反馈型。

在算法生成型中，Critic模块的奖励信号由算法根据任务目标自动生成。2012年Mahmoudi等人[62-63]将Actor-Critic算法应用于BCI，提出了一种基于Hebbian强化学习的自适应神经假体控制器，结果表明该控制器能够在模拟环境中快速收敛到有效的控制策略。在被试反馈型中，Critic模块的奖励信号直接来源于被试的反馈，例如特定脑区神经信号，这种方式更贴近实际BCI应用，能够根据被试的真实意图进行学习。2011年，Mahmoudi等人[64]提出了一个Actor-Critic框架，利用伏隔核（Nucleus Accumbens, NAcc）的神经活动作为Critic模块的输入，提供直接来自大脑的评估反馈，从而实现对BCI解码器的自适应调整，结果表明该方法能够在模拟环境和闭环实验中维持高性能的神经命令解码。2020年，Shen等人提出了利用内侧前额叶皮层（mPFC）的信号作为Critic来辅助多步骤解码任务的学习，结果显示与仅使用最终奖励的模型相比，上述方法平均收敛速度更快，稳定性更高。

1.3 个性化强化学习与基于迁移学习的强化学习

随着BCI任务复杂度和用户数量的增加，个性化学习与迁移学习成为满足多样化需求的关键方向。不同用户的运动习惯和神经信号模式各异，因此需要更具针对性的解码方法。

逆向强化学习通过从用户行为中推断奖励函数，实现个性化的意图建模。Tan等人[65]提出了一种基于隐藏脑状态的核逆向强化学习方法，通过隐藏脑状态估计和核方法来实现个性化解码，在大鼠的实验数据上得到了验证，该方法推断出的状态值分布与行为轨迹高度一致，比手动设计的奖励信号更准确地反映被试的内部奖励评价。

在使用BCI的系统时，每个用户的集中度（任务参与度）也存在差异，这直接影响BCI系统的训练效率和解码性能。Zhang等人[66]提出了一种在线强化学习解码方法，通过建模任务参与度调节解码器的学习速率，从而优化训练过程，研究表明该方法在大鼠的脑控任务中，与未建模任务参与度的解码器相比，能够使用更少训练数据达到相同的收敛性能，同时显著提升解码准确性和轨迹重建效率。

迁移学习旨在利用从一个任务或领域中学到的知识来加速对新任务或领域的学习。迁移学习本身并非一种独立的解码算法，而是提升解码器泛化能力的机器学习策略，通过迁移已有任务中的神经模式知识，辅助目标任务中解码器的训练过程，从而缓解训练数据不足和跨任务适应性差的问题。在脑机接口中，迁移学习常与解码算法（如CNN、RL）结合使用，共同实现更高效的跨个体、跨任务神经解码。由于迁移学习可以帮助解码器快速适应新的任务或跨会话的变化，而无需从头开始训练，显得尤为重要。Zhang等人[67-68]提出了一种基于核强化学习的聚类迁移学习方法，用于提升跨任务适应性，通过对历史任务中的神经模式进行聚类，提取神经信号到动作的条件分布，并将最相似的聚类结果应用于新任务的神经模式，解码器根据已经训练好的参数来加速新任务的学习过程，从而提高了系统的泛化能力。

目前，RL在BCI中的研究已从基础算法探索转向技术优化。基于Q-value的算法通过单步和多步解码改进提升性能；Actor-Critic算法则通过策略优化和用户反馈增强适应性。个性化强化学习与基于迁移学习的强化学习将是未来应对复杂任务和多样化需求的关键。

- 基于不同空间时间尺度的解码

为了解析大脑中神经网络在各个空间尺度和时间尺度的编码原理和学习机制，当前点过程模型围绕单神经元编码、群体计算原理和学习适应性机制三大方向展开。

2.1 单神经元编码模型构建

单神经元编码模型的核心目标为建立输入（感官刺激、运动意图等）-输出（spike序列）的映射关系。此类模型使用调谐函数评估神经元应对特定输入特征的偏好，并使用随机过程模拟峰电位发放的随机性。例如广义线性模型（Generalized Linear Model, GLM）在2004年被应用于调谐函数将高维度的视觉刺激投影到标量化的峰位发放概率，揭示了初级视觉皮层神经元对方向选择性的量化规律[69]。而针对时间动态特性更明显的运动区神经元，非同质泊松过程的瞬时线性-非线性-柏松模型（instantaneous-LNP）实现了神经元对于连续动作的一对一实时编码，并发现模型参数可解释运动皮层对位置、速度的偏好性编码[70]。2022年，Chen等人在instantaneous-LNP模型基础上揭示了通用调谐函数参数—神经调制状态，具有有限范围的特点，有助于学习过程中突变参数的追踪[71]。

2.2 神经群体计算原理

核心问题在于解析神经元群体如何通过协作实现稳定编码，研究主要有三种路径：1）相关性分离：该路径基于单神经元的调谐函数，将神经元的调制分解成独立编码和合作编码两部分。其中独立编码的部分仍旧和目标状态直接相关，而合作编码部分有不同的表现形式。例如Kim等人使用格兰杰因果推断构建了神经元间因果连接，并用神经元交互部分表征合作编码部分[72-73]；Sani等人发展了优先子空间标识（preferential subspace identification, PSID）将神经元调制分解成目标状态相关和目标状态无关两部分，揭示了与目标状态相关的神经动力学[74]。然而，这些方法提高了神经编码和解码的准确性，但合作编码的细节仍旧模糊。Li等人通过结合独立编码模型和神经交互模型发展了神经相关性点过程滤波器（correlation integrated point process filter，CIPPF），首次将神经解码在点过程观测中分解成独立编码和合作编码，提供了更清晰的神经群体编码框架[75]。2）低维流型理论：该路径通过降维简化高维信号，使用低维空间表征神经群体连接以及神经群体信号的子空间。2020年，Gallego等人通过主成分分析法建立了神经元活动的潜在低维空间，发现低维空间中的编码更加稳定是神经活动维持稳定编码的潜在解释[76]；2023年，Abbaspourazad等人在构建低维流形时，添加了神经活动状态的动力学模型，能够更好的预测和解码神经元活动[77]。3）递归神经网络拟合：基于递归神经网络训练出一个具有外部输入，神经元全连接的网络结构，能灵活模拟群体动态。DePasquale等人通过全连接的神经网络学习了神经元群体的神经动力学，并使用外部输入模拟神经系统与环境的交互成功调制了神经群体的编码[78]，但该方法得到的网络结构不具备唯一性，无法正确表征实际的神经元连接带来的协作关系。

2.3 神经元学习适应机制

神经元学习适应机制通过探究神经元调谐特性的变化与神经相互作用强度的调整，揭示大脑如何通过更高效的编码优化任务记忆功能——既调整单元特性，又重塑功能性连接。在现有研究中，通过追踪单神经元模型中调谐函数参数的变化，单个神经元独立编码在快速学习。Wang等人使用双重蒙特卡洛点过程估计追踪了单神经元编码函数中调制深度的缓慢变化[79]。Chen等人通过全局追踪有限范围的神经调制状态，成功追踪了学习中快速变化的神经元偏好方向。在神经群体活动中，神经元功能性连接的适应性也得到了一定程度的揭示。Yu等人在脑机接口的脑控实验中发现，轻微扰动下，神经群体往往在原有的低维流型中改变神经元活动来适应新的脑控系统，但在严重干扰中神经系统会发展出新的低维流型以抵抗扰。Chen将Adam应用于点过程模型，直接追踪到了神经元连接强度在学习过程中的变化[80]。然而，这些模型中独立编码和合作编码在适应性学习中的关系仍旧模糊。

- 基于跨脑区通信的解码

随着闭环调控系统与神经解码算法的发展，跨脑区神经旁路系统研究已取得突破性进展，主要梳理两大研究方向包括基于神经旁路系统在不同神经通路上的应用和针对神经信息传递的跨脑区预测模型的发展。

3.1 神经旁路系统的研究进展

针对脊髓损伤所致瘫痪患者的运动功能重建，当前研究发现可通过解码初级运动皮层神经电信号获取运动意图，并利用动态微电流刺激技术作用于特定肌群，实现上肢部分运动功能恢复。然而，此类基于外周肌肉刺激的方案易引发疲劳效应，限制了其长期日常使用。Grégoire Courtine团队提出了结合运动解码与脊髓硬膜外电刺激的方案，使下肢瘫痪患者恢复部分下肢功能[81-83]。这类基于人工神经旁路的运动康复方案突破性在于直接利用患者自身神经系统和肢体实现运动输出，相较于传统脑机接口中机械臂等外设控制方案更具生理相容性。

在认知功能领域，神经旁路系统同样展现出重要价值。阿尔茨海默症引起的海马体内神经突触异常会导致记忆衰退，为此研究者开发了记忆神经假体，通过在海马体CA3区施加动态闭环的多电极微电流刺激，成功改善了多个被试在图像记忆任务中的准确率[84-85]。相关记忆假体技术路线有两类。其一采用类似运动神经旁路的方案，通过解码海马体CA1区（上游）神经活动识别记忆内容的类别，进而向CA3区（下游）施加特定刺激模式[86]。其二则使用基于机器学习的跨脑区神经脉冲预测模型，直接从CA1区神经活动预测CA3区神经脉冲序列，并将其用于指导电刺激的时空间模式[87]。

3.2 跨脑区神经脉冲预测模型

神经旁路系统的核心挑战在于如何依据上有脑区记录的神经电信号，实时生成符合生物编码规律的下游微电流刺激的动态模式。基于跨脑区神经脉冲预测模型构建的刺激范式，因其能够模拟生理神经信息的传递机制，有望在功能康复效果上展现显著优势。

早期研究基于点过程模型框架开展。广义线性模型通过优化泊松似然函数，实现了对神经元脉冲发放概率的建模[88]。其随后被拓展为广义 Laguerre-Volterra模型（Generalized Laguerre-Volterra Model，GLVM），引入Volterra核函数，整合上游神经脉冲间的非线性相关性，该模型已成功应用于前述记忆神经假体。基于人工神经网络的点过程模型，通过捕捉神经元集群间的复杂非线性关系，进一步提升了模型预测准确率[89]。

然而，传统点过程框架下优化似然函数的方法属监督学习，需依赖完整的下游脑区信号记录。对于神经通路损伤患者，下游脑区记录难以获得，局限了这类模型的临床应用。Wu等人[90]于2023年开创性地提出强化学习驱动的跨脑区神经脉冲生成模型，该方法使用行为层级的奖励信号优化模型，针对性地优化模型生成的神经脉冲序列在行为任务上的解码效果，摆脱了对下游神经记录的依赖，取得了较监督学习方法更优的解码成功率。后续改进方法通过引入神经流形约束，令强化学习驱动的跨脑区模型考虑神经元集群的协同放电规律[91]。该方法生成的神经信号在保持行为解码表现的同时，更严格遵守了神经元放电的编码性质，有望在实际应用电刺激时取得更安全、稳定的临床效果。

# （二）非侵入式脑机接口软件

受限于非侵入式脑信号天然存在的低信噪比与高个体差异，信号处理与解码算法的研究成为非侵入式脑机接口领域的核心问题之一。传统的脑信号处理主要依赖人工设计的特征提取与经典机器学习算法进行分类与识别，包括空间滤波（如共空间模式（Common Spatial Patterns，CSP）、扩展DAWN算法（xDAWN）、时频分析、小波变换、独立成分分析（Independent Component Analysis，ICA）等方法，在标准范式和受控环境下取得了一定效果。CSP是一种经典的空间滤波算法，主要应用于MI脑机接口任务中。xDAWN算法可增强与目标刺激相关的ERP成分，提高信噪比，并有效提取出与刺激事件相关的脑电活动。时频分析可提取出脑信号中关键的时频特征，小波变换则用于实现脑电信号的频段分解。ICA 是盲源分离算法，将混合信号分解为一组统计上彼此独立的成分，常用于去除眼电、心电与肌电干扰等伪迹。

然而，随着应用场景的复杂化和实时交互需求的提升，传统方法逐渐暴露出适应性差、泛化能力差、特征表达能力有限等问题。近年来，深度学习技术的快速发展为脑信号的自动特征学习和复杂模式建模提供了新的解决方案[92]。

卷积神经网络（Convolutional Neural Network，CNN）能够有效提取脑电信号的空间与时域特征，适用于EEG信号的时空建模；循环神经网络（Recurrent Neural Network，RNN）和长短时记忆网络（Long Short-Term Memory，LSTM）建模脑信号的时序依赖，适合连续任务；图神经网络（Graph Neural Network，GNN）建模EEG通道的图结构，能够捕捉脑区之间的复杂连接关系；Transformer可高效建模非侵入式脑信号的长程依赖，有效捕捉上下文关系，为脑信号建模提供了新的技术路径[93]。针对非侵入式BCI普遍面临个体差异显著、跨设备/跨任务迁移困难等问题，当前已有多种解决方案。

- 脑电基础模型预训练

自监督学习与大模型预训练技术已开始应用于脑信号的特征学习，通过构建大规模的脑信号预训练模型，提升了小样本场景下的解码性能[94]。研究者开始面向海量脑电乃至跨模态生理信号构建脑电基础模型，显著缓解了“小样本标注不足、个体差异巨大”的长期瓶颈。LaBraM提出基于Transformer的EEG大规模通用预训练模型。EEGPT[95]通过自监督预训练捕获时序脑电的全局特征。Banville等[96]率先将多种自监督预任务应用于临床EEG，验证了无监督特征对异常检测任务的可迁移性。BENDR[97]基于Transformer，并引入对比自监督任务，在百万级 EEG 片段上预训练通用表征。MAEEG[98]采用“遮蔽-重建”策略的Masked Auto-Encoder，在大规模 EEG 上学习局部—全局一致特征。	BrainBERT[99]针对颅内SEEG设计掩码频谱建模，学习复杂非线性转换。Brant系列（Brant[100]、Brant-2[101]、Brant-X[102]）依次实现了长程依赖建模、头皮-颅内信号融合及跨生理信号对齐。BIOT[103]建立跨模态、跨任务生理信号联合预训练框架。EEG2Rep[104]在潜在空间预测被遮蔽片段，强化时序一致性和判别能力。

总体来看，“大规模数据+自监督预训练+轻量微调”已成为当前脑电基础模型的主流范式，为多任务快速迁移及高精度实时脑电解码奠定了算法基础。

- 迁移学习与域适应算法研究

迁移学习与域适配算法利用其他辅助数据或知识来加速目标用户校准，可提升非侵入式脑机接口跨被试、跨设备、跨任务的泛化能力提供了技术支撑[105]，有效地解决个体差异带来的性能瓶颈问题。迁移学习被认为是能处理个体差异、减少被试校准时长的极具前景的方法，方法利用其他辅助数据来帮助目标用户的精准解码。当前脑电解码中的的迁移学习研究聚焦于离线及在线场景：离线场景即使用全部的目标用户数据来设计算法缓解跨域分布差异；在线场景即目标用户的数据在线到达，不能一次性使用。针对EEG信号用户个体差异大的问题，He和Wu[106]提出高维空间欧氏对齐算法，显著降低了源用户和目标用户的数据分布差异。对抗防御算法可有效提升模型对信号扰动的鲁棒性，进一步保证脑机接口系统的安全性。Li等[107]提出针对在线实时BCI系统的迁移学习算法，有效提高了在线EEG解码的准确性。

- 隐私保护算法研究

随着非侵入式脑机接口的应用场景逐步涉及个体健康管理、消费电子等领域，脑信号数据的隐私安全成为必须正视的问题。现阶段，BCI系统的隐私信息包括数据隐私和模型隐私，数据隐私如用户的身份、性别和认知等，模型隐私如模型架构、模型参数等[108]。近年来社会各界对个人隐私的关注日益增加，一些法律法规相继制定，如我国《个人信息保护法》、欧盟《通用数据保护条例》等关于生理信息保护的条例，因此，用户的隐私保护问题亟待解决。多种策略如联邦学习框架、差分隐私机制和脑电数据生成等均可提高BCI系统的隐私保护能力。联邦学习框架支持多中心协同训练，避免原始脑电数据泄露。差分隐私机制对合成模型或合成数据添加噪声，对于查询任务加入一定的偶然性，使外部攻击者难以通过分析数据或模型结果推测个体隐私。具体实现中，可视隐私需求与系统设计采用拉普拉斯或高斯噪声。该策略在保持模型性能的同时，防止敏感数据泄露，增强参与用户的隐私保障。其还可通过调整隐私预算，在隐私保护与模型精度之间取得平衡，适应不同场景需求。脑电数据生成通过合成新的脑电数据进行共享，保护源数据隐私。数据去识别化与匿名化构成另一条重要的隐私保护路径。去识别化是指在处理敏感数据时，首先彻底剔除可直接指向个人身份的字段（如姓名、性别等），从源头上阻断数据与特定个体的关联。匿名化则将剩余信息转换为不可逆的形式，通过脱敏或替换等手段，使任何后续分析者均无法反推个人身份。将这两种技术引入联邦学习，可在“看不到原始数据”的前提下实现高效聚合与模型训练，不仅提升了整个信息处理环节的安全性，也增强了联邦系统的可信度，同时最大限度地减轻对任务性能的潜在影响。

# 三、脑机接口外设

# 四、脑机接口临床应用

# （一）侵入式脑机接口临床应用

侵入式脑机接口应用主要面向医疗领域，包括运动功能重建、感知觉增强、脑疾病治疗等方面，国内外已分别实现了诸多应用成果，欧洲和美国等国家和地区仍处于领跑主导位置。

- 运动脑机接口

全球人口迅速老龄化，罹患瘫痪、退行性神经疾病或精神疾病群体数量剧增，脑机接口在帮助运动障碍病患恢复运动功能方面有广阔前景。

瑞士洛桑联邦理工学院Micera团队开发了一种有同步神经内触觉和位置反馈的闭环神经假肢，通过在上臂的正中神经和尺神经植入横向束内多通道电极（TIME）进行神经内刺激作为感觉替代，提供位置信息和躯体定位的触觉反馈，成功使前臂截肢的被试恢复了接近自然的本体感觉敏锐度，能通过握持辨别物体的大小[109]。美国巴特尔纪念研究所Ganzer团队实现一位完全性脊髓损伤的患者利用自身手部残留的触觉信号，恢复运动功能和触觉感知，从而实现由触觉信号调节的抓握动作恢复[110]。美国Neuralink的柔性脑机接口设备已成功将临床试验扩展至海外，全球范围内开展多例临床试验，除了基本的意念控制电脑、玩游戏，患者现已能通过意念操控机械臂完成更复杂的动作。Synchron公布了65岁的ALS患者Rodney Gorham使用其血管介入式脑机接口的5年经历，不仅未出现与设备相关的严重不良事件，其控制能力也从最初的“单击”进化到“2D光标控制”，并能通过意念操控家中的智能音箱、风扇、灯光、宠物喂食器等设备，展示出系统的长期稳定性和可进化性。

国内博睿康自主研发的植入式脑机接口手部运动功能代偿系统创新产品NEO已获国家药监局批准上市，标志着国际首个侵入式脑机接口医疗器械进入临床应用阶段，该产品完成了共36例颈段脊髓损伤截瘫患者的植入手术，所有患者均成功实现居家脑控抓握辅助与康复训练，在主要临床终点上100%达标，且未见医疗器械相关严重不良事件，初步验证了硬膜外微创植入路径的安全性和长期使用可行性。阶梯医疗自主研发的256通道无线高通量侵入式脑机接口系统（WRS02）已完成3例临床植入，并成功验证脑控交互功能，是目前已进入国家药监局创新医疗器械"绿色通道"的侵入式脑机接口产品，计划2026年中启动大规模多中心注册临床试验，年内完成约40例患者入组。脑虎科技的“植入式脑机接口手部运动功能代偿系统”（即“三全”系统：全植入、全无线、全功能）正式进入国家药监局 CMDE 创新医疗器械特别审查程序 ，是国内首个获此资格的硬膜下植入式柔性脑机接口产品。浙江大学牵头完成全国首例闭环脊髓神经接口植入手术，帮助患者逐渐恢复下肢运动功能，让其可以直线行走，还实现转弯、上下坡、变速行走等多种日常活动，经过几个月训练，可以脱离助行器、拄拐自主行走。团队随后又陆续为六位患者进行了手术，为脊髓损伤患者提供了全新的功能恢复路径，在运动、感知及生活自理能力上带来了突破性的改善。南开大学段峰教授团队完成了中国首例介入式脑机接口辅助人体患肢运动功能修复试验，系统通过颈部血管微创手术，将直径2mm的电极支架送入大脑运动皮层血管，并同步植入无线供能与传输装置，建立起"脑—机—肢体"闭环通路。术后两周内，患者即可意念控制手部抓握，信号传输稳定率达99.7%，续航5年以上。团队已启动多中心临床试验，并在北京、武汉等地建立专病门诊。

- 语言脑机接口

言语交流是人类日常生活中最基本的社交需求之一，神经系统疾病或创伤（如中风、渐冻症、脑干损伤等）可能导致患者完全或部分丧失言语能力，严重影响患者的生活质量与社交参与能力[111-113]。通过绕过受损的神经通路，直接从大脑活动中解码患者的交流意图，脑机接口技术为解决此类交流障碍提供了一条全新的途径。

2.1 基于颅内脑电技术的言语产生神经编码机制研究

早在上世纪中叶，神经外科医生如Penfield等人就已通过术中电刺激定位的方式确定了腹侧感觉运动皮层对于言语构音器官（如唇、舌、下颌、喉部等）的控制作用[114]，但具体神经编码机制尚未得到明确揭示。2010年以来，随着更高空间和时间分辨率的颅内脑电技术的广泛应用[115]，研究人员首次在人体内清晰地揭示了言语产生过程中构音运动的精细神经表征。

言语产生涉及音段特征（如音素、元音与辅音的产生）与超音段特征（如音调、韵律、语调）的协调控制[116-118]。大量研究表明，这两类特征的神经编码机制存在明显的空间分工与功能分化。音段特征的神经编码主要集中在中央前回腹侧区域（ventral precentral gyrus），以不同构音器官运动在空间上特异的神经激活图谱为特征。研究人员发现在言语相关构音运动发生时，腹侧感觉运动皮层（ventral sensorimotor cortex，vSMC），尤其是中央前回的腹侧部区域（ventral precentral gyrus），表现出明显且特异的高频Gamma（70–150Hz）能量增加[119]，这种神经信号的变化与特定的构音器官（如唇、舌、下颌、喉部）的运动轨迹和模式紧密关联，这一区域也被称为言语运动皮层（speech motor cortex）。2013年，Bouchard等[120] 通过ECoG技术首次绘制出言语产生时不同构音器官运动在皮层表面的详细时空动态分布图谱。 2018年Chartier等[121]进一步揭示了ECoG电极记录到的言语运动皮层的局部神经活动编码表达了不同构音器官组合的动态构音运动轨迹（articulatory kinematic trajectory, AKT）。这些对于音段特征神经编码机制的研究发现，奠定了言语产生神经编码机制研究的重要基础。

相对而言，超音段特征则与喉部、声带运动和音高控制相关，涉及不同的大脑区域与功能模式。近年来，研究人员也系统地探讨了超音段特征所对应的喉部运动与音高控制在不同语言类型中的神经基础与功能意义。Dichter等人[122]通过ECoG技术首次揭示了背侧喉部运动皮层（dorsal laryngeal motor cortex）在非声调语言（如英语）的语调（intonation）控制中发挥重要作用，其研究发现该区域神经活动与语调变化紧密相关，特别是在疑问语气和强调表达的情况下，这些神经活动呈现显著的调制变化。

作为典型声调语言的汉语，其言语产生过程涉及更加复杂的音高控制机制，需在音高的正向（升调）与负向（降调）两个方向进行精确的实时调整。复旦大学附属华山医院和上海科技大学团队[123-124]通过对中文使用者进行ECoG记录研究，进一步揭示了声调语言特有的音高控制神经机制，发现双侧背侧喉部运动皮层（bilateral dorsal laryngeal motor cortex）以及左侧腹侧喉部运动皮层（left ventral laryngeal motor cortex）共同参与了汉语语音音高的精细控制。研究明确指出，这些脑区在汉语声调产生过程中表现出显著且特异的正向或负向神经调控模式，这种模式不同于非声调语言的单向语调调控机制，更能体现声调语言独特的神经基础。

综上，不同语言类型中音段特征与超音段特征的神经编码机制存在明显差异。音段特征编码主要体现为中央前回腹侧区域不同构音器官的空间拓扑分布，而超音段特征则更多涉及喉部与声带运动的精确时序与音高调控机制。这种音高控制机制在声调语言与非声调语言中具有截然不同的功能和神经表征特征，尤其在汉语这样的声调语言中更为复杂和精细，这为后续语言脑机接口系统开发提出了更为丰富的神经生理学挑战与研究机遇。

大多数语言产出机制研究聚焦于发音控制，语言产出的另一模态是书写。汉字的书写难度远高于英语等字母文字的书写难度，单个汉字包含字、部首、组件、笔画等多层级的书写单元。浙江大学团队[125]发现手结区MEA编码汉字书写的单元是比笔画更小的笔画片段，在每个笔画片段之内，手结区神经元通过近似线性的形式编码运动速率，但不同笔画片段所依赖的神经编码机制不同。这表明，手结区可能利用了一种多层级的方式编码汉字书写，大脑首先将汉字字形拆分为笔画片段，然后在每个笔画片段的过程中维持稳态、线性的神经编码，在笔画片段切换时同步切换神经编码方式。汉字书写是涉及字形提取，字形分解，运动执行等的复杂过程，这一过程涉及脑区不止手结区一个区域，但是其它脑区在汉字书写过程中的编码形式仍待探索。

2.2 基于颅内脑电技术的离线语言解码技术研究

在言语产生神经编码机制明确后，研究人员进一步探索如何从这些神经活动中解码言语内容，以实现基于脑机接口的沟通恢复。这一研究阶段通常采用离线解码（offline decoding）的方式进行，即首先记录参与者的神经数据，再通过机器学习算法分析这些数据，以评估解码特定言语单元（如音素、音节、单词或句子）的准确性与可靠性。

离线解码的典型实验方案为指示参与者在明确的提示下发出特定的词语、短语或声音（如辅音-元音对“ba”等），然后记录此过程中的颅内脑电（如ECoG或MEA）[126-127]，随后这些记录的神经数据被分段、对齐至目标言语事件（例如词语发音起始点），并通过各种机器学习算法进行分类或重构，通常涉及两种离散式解码（discrete decoding）和连续式解码（continuous decoding）主要解码任务。

在离散式解码中，算法的任务是在一个有限的词汇集中识别发声的具体词汇或音节[128-133]。这些研究采用了一系列统计学习方法，包括模板匹配、线性判别分析（LDA）、逻辑回归、支持向量机（SVM）以及近年来广泛使用的深度学习模型。早期研究表明10个词汇分类任务的准确率约为48%；到2020年左右，利用ECoG数据进行的57个辅音-元音音节的分类准确率提高到55%[134]，而MEA数据用于39个音素分类的准确率达到了33%。

相比之下，连续式解码则更加复杂。其目标是直接重构连续的言语声音波形，而非仅对有限的分类标签进行识别。连续解码的输出空间更为庞大，但理论上能提供更灵活、表达力更丰富的沟通方式，并能捕捉言语的超音段特征（如音高、响度和节奏）[135]。目前，这类研究使用的解码方法包括线性回归、单元选择（unit selection）方法和深度学习模型（如长短期记忆网络LSTM）[136]。通常，这些研究首先将神经活动映射到声学特征（如频谱[137]）或构音运动特征（如构音器官运动轨迹AKT，然后再通过声码器算法（如WaveNet、WaveGAN或LPCNet）将这些特征转换为实际的声音波形。加州大学旧金山分校（UCSF）团队（Anumanchipalli等[138]）提出一种双阶段深度学习解码方案：首先使用LSTM网络从ECoG记录数据解码出构音器官运动轨迹，再利用另一个神经网络将构音运动特征转化为声学特征，最后重建言语波形。虽然重构的语音整体上仍不完全可理解，但已能成功区分发声与静音状态，并初步呈现原始语音的频谱轮廓。后续研究表明，利用更大规模的数据训练、更高密度的电极阵列以及更复杂的神经网络结构，连续言语重建的质量逐步提高，部分句子在特定上下文中甚至可以实现较好的可理解性。

近年来以汉语为代表的声调语言解码研究也取得了突破，积极探索了解码技术层面如何通过先进的深度学习模型处理汉语音高的连续动态特征。上海科技大学、华山医院等团队利用ECoG和立体定位脑电（Stereoelectroencephalography, SEEG）信号，成功实现了中文单字、词语及连续句子的离线解码[139-140]，结合深度神经网络，可以有效地捕捉汉语声调的复杂时频动态特征，并在高频Gamma神经活动中精准解码出目标语音内容与音调信息并合成输出汉字语音以及语句文本。脑虎科技联合复旦大学附属华山医院等团队展示了一套能够实时解码中文全音节的脑机接口系统，系统覆盖了现代汉语中所有394个基本音节，通信速度接近50字/分钟，逼近正常人手写输入的速度[141]。上海科技大学联合复旦大学附属华山医院搭建了汉语语句的解码算法架构，将汉语脑机接口从“字”拓展至“句”的层面，首次实现从“意念”到汉语语句文本的解码。单受试者最高声调解码正确率达到95%，单字解码准确率达88%，字误率仅为14%。5名不同年龄、性别、教育程度、方言背景的受试者的平均声调分类准确度亦达93%，单字解码准确率达82%，字误率为21%，证明该算法模型具有较高的跨被试稳定性与可推广性。同时，团队还实现了“耳语音” （whisper）状态下的汉语句子解码，进一步验证了该模型未来用于帮助失语患者重建汉语语句表达的可行性[139]。

离线解码研究也存在其挑战和限制，健康参与者在发声时伴随的体感觉、本体感觉和听觉反馈，在实际BCI用户中往往缺失或严重受限，这导致了神经活动的差异。此外，言语发声过程中的物理运动可能导致记录信号中的机械干扰，这一干扰需要通过严格的实验设计与信号处理技术进行有效的区分与消除[142]。

2.3 基于颅内脑电技术的实时语言解码技术研究

实时在线语言脑机接口要求系统实时处理和解码脑活动信号，立即将解码结果反馈给用户，从而形成完整的闭环交互系统。

最初在线语言脑机接口研究的词汇量较小，重点在于如何证明实时解码的可行性。研究人员在植入ECoG电极的患者中开展实时的语音解码实验，解码任务仅局限于几个单词或元音类别，通过在线神经信号解码[143]，可实时识别9个可能的问题与24个可能的回答，系统利用预先的语音感知-生成混合解码方案完成了单词级的实时解码。此后，词汇量和解码的准确率逐步提升。加州大学旧金山分校团队首次实现了实用性较强的实时在线言语解码[144]，因脑干中风而失去发声能力患者通过植入128通道ECoG电极阵列，能以98%准确率实时检测用户尝试发声的动作，从50个单词组成的词汇表中以74%的准确率实时在线解码用户试图表达的具体词汇，词汇规模扩大至超过1000个可实现英语和西班牙语双语输出，极大提高了实时在线言语解码系统的临床实用性[145-146]。2023年，斯坦福大学等高校的两项临床试验研究进一步显著提升了解码的准确率与实用性。Willett等人[147]在渐冻症患者上开展临床试验，通过中央前回和额下回植入的4个64通道微电极阵列记录神经活动，利用深度学习算法解码合成文本，在50词语料库中实现了91%的解码准确率，并在125000词语料库中实现了74%准确率，达到了62词/分钟的解码速率。Metzger等人[148]则进一步拓展了语言脑机接口的功能，将用户尝试发声时的神经信号解码为语音，同时驱动虚拟面部动画同步展示相应的口型和面部表情。利用系统在因脑干中风导致说话能力受损的患者中植入253通道ECoG电极，成功在1024单词词库上实现了78个英文单词/分钟的解码速率，达到75%解码正确率。Kunz等人提出“内心言语”数据训练人工智能模型，可从包含12.5万个单词的词汇库中解码想象中的句子，准确率达74%，并能捕捉用户默念的数字，为帮助严重言语障碍患者重建自然交流能力开辟了新路径[149]。加州大学戴维斯分校团队[150]通过增加电极数目至256个提升了系统性能，仅通过30分钟的训练数据，便可在50个词汇的词汇表中实现了99.6%的解码准确率，次日词汇表扩大到125,000个词汇，经过额外1.4小时的训练数据，可达90.2%准确率，八个月后仍保持在97.5%以上，证明了实时在线语言脑机接口的实际应用潜力。

此外，研究人员还积极探索了连续语音合成技术。2023年，系统能够合成出具有一定可理解性的语音片段，用户从1024个单词的词汇集中合成出的语音，人类听者能以46%的准确率正确转录[148]。基于非因果模型，仅能在语句完成后利用完整的序列进行语音合成，未达到完全实时的连续解码与播放，实现灵活、即时且持续可理解的实时在线语音合成接口仍是当前领域的重要挑战之一。2025年，最新研究通过ECoG电极和深度循环网络实现了仅有80ms延迟的即时语音合成[151]。西湖灵犀科技杨杰等人提出了全谱系中文语言解码框架，每秒可完成超过 200 次解码运算，端到端延迟低于10ms，能实现每分钟30个汉字的实时脑电解码，准确率达70%以上，为失语症患者恢复自然语言交流能力开辟了新路径。

- 视觉脑机接口

神经工程、仿生学、微电子、计算和无线技术等相关技术的进步，推动视觉脑机接口，尤其是皮层视觉脑机接口快速发展

3.1 视网膜假体

视网膜假体原理是利用患者尚且保留的视觉通路（视神经与视觉皮层功能完好），通过植入微型电极阵列对残存的视网膜神经元进行刺激，以实现外部光信号到电信号的转换[152]。这种方法通常适用于患有视网膜退行性疾病，如色素性视网膜炎（Retinitis Pigmentosa, RP）和老年性黄斑变性（Age-related Macular Degeneration, AMD）等而导致失明的患者[153]。与其它刺激方式相比，视网膜假体可以利用患者自身部分保留的视网膜结构（如双极细胞、神经节细胞等），对光信号进行初步处理后再传递给大脑。根据植入层次的不同，视网膜假体可大体分为视网膜上假体（epi-retinal prosthesis）、视网膜下假体（sub-retinal prosthesis）和脉络膜上假体（suprachoroidal prosthesis）等路线[154]。视网膜假体已被开发为治疗视网膜病变导致失明的潜在治疗方案。

（1）视网膜上假体（Epiretinal Prosthesis）

这类假体通常贴附于玻璃体腔内侧的视网膜表面，并通过外部获取的图像信号来驱动电极刺激视网膜神经节细胞（RGCs）[155]。其优势在于植入手术相对熟悉，可在较宽的视野范围内布设电极，从而在一定程度上帮助患者进行环境定向和物体轮廓识别[156]。然而，上视网膜假体往往需要通过视网膜钉进行固定，长期稳定性以及电极与神经元的有效接触仍是挑战。此外，由于神经纤维层的存在，轴突激活导致的光幻视呈条纹状或拉伸态，也会影响到视觉质量。

Argus II是最早获得美国FDA批准并进行商业化应用的上视网膜假体系统[157]。Argus II视网膜假体有60电极阵列，目前已植入超过350名患者。Argus II的组成部分包括了安装在眼镜上的摄像头，用以捕获环境中的视觉信息，这些视觉信息由视觉信息处理单元（Visual Processing Unit, VPU）进行处理，并通过外部射频发射器（external radio frequency transmitter）无线发送至植入物接收器及专用集成电路中。Argus II的电极通过视网膜钉固定在视网膜上，其通过永久性的巩膜切口连接到集成电路上，通过电极刺激视网膜神经节细胞。植入该假体的患者获得了视觉功能的改善，Breanne等人使用单字母识别任务测试患者对字母识别的能力，其中两名被试的识别正确率可以达到75%[158]。Beyeler 等人使用单字母识别测试来评估4名植入假体的患者的视觉功能改善情况，其中患者植入后的视力范围为20/15887至20/1260，但所有患者均出现视力缺陷，导致错误识别及识别时间增加[159]。目前植入Argus II假体的患者最高视力为20/1260，恢复的视野最大范围约为11deg×19deg。此外在使用电极的顺序刺激时，植入Argus II的患者对正方形大小辨别任务和简单形状辨别任务的表现更好，但在更复杂的辨别任务中，顺序刺激并没有表现出明显改善。目前，Argus II已因企业战略转向等原因停产，但其开创性的临床应用意义仍然重大。

Pixium Vision SA开发的视网膜上假体IRIS，该假体具有150个电极组成的刺激阵列，阵列连接到柔性的眼内箔片上，箔片连接到用于无线传输和产生刺激电流的电子设备，用以刺激残留的视网膜。在2016至2017年，在10名RP患者中植入IRIS II视网膜上假体[160]。患者对运动方向的检测等视觉功能的测试结果有所改善，方块刺激定位测试的平均误差距离从8降低到2，图像识别的准确度从45%提高到了55%。不过由于生产和配套问题，该产品也已退出市场。IMI 视网膜上假体使用直径为0.10-0.36mm的电极，使用聚酰亚胺的微接触膜作为基底。在玻切手术后，将微接触膜和电极放置在患者的玻璃体腔内，通过电刺激激活视网膜，共在20名视力范围从20/1000至无光感的被试中进行植入。植入后3个月无严重不良事件发生，且几乎所有受试者在采访时都描述了几种具有不同特征的视觉感知[161]。国内湘雅医院于2026年4月完成了微智IMIE智能视网膜的植入手术，暖芯迦也研发了E-BCI视网膜上假体等视觉修复技术。

（2）视网膜下假体（Sub-retinal Prosthesis）

视网膜下假体植入于视网膜色素上皮与神经感受层之间，能够更好地利用双极细胞进行初步视觉处理，从而有望获得更细腻的图像分辨率[162]。同时，此种植入方式通常无需使用视网膜钉固定，机械稳定性更优。但其手术相对复杂，若要实现大范围刺激也更具挑战。

视网膜下植入物中，最先获得CE批准的是德国的Retinal Implant AG公司的Alpha IMS，后推出了其改进的版本Alpha AMS。Alpha AMS假体的视网膜下植入部分的尺寸为4.0×3.2mm，包含了一个2.8×2.8mm的电极阵列，相当于9.3deg的视野范围。电极阵列包含了1600个像素，其中每一像素的尺寸为70×70μm，包括光电二极管、一个放大器和一个刺激电极。Alpha AMS假体植入视网膜下与色素上皮之间，通过刺激残留的双极细胞等，恢复患者的视觉感知。Alpha AMS视网膜下假体无外置眼镜，使利用眼睛的自主运动直接激活局部视网膜成为了可能。然而，由于环境光强度水平不足以为电刺激提供足够的电力，因此需要通过经巩膜的导线，将视网膜下芯片连接到眼外电子元件，该电子元件本身通过无线连接到外部电源。在植入Alpha AMS假体的患者视觉功能均能够得到一定的改善，Katarina等人报道了3名使用Alpha AMS视网膜下假体的患者，能够在26个字母识别测试中识别出5个字母[163]。在植入的病人中，有6名患者可以通过Landolt 视力表测试视力，视力范围从20/2000至20/546[164-165]。最近关于Alpha AMS假体的加速老化寿命测量结果显示，该植入物的临床寿命约为5年[166]。然而，由于假体制造与植入成本高昂，加之公司于2019年解散，目前该视觉假体的量产和推广均已中断。

以色列的Nano Retina公司开发的NR600视网膜下假体也进入临床试验阶段。NR600是一款独立的视网膜下假体，无导线等连接，无需跨眼电路，需要搭配特制的无线可充电眼睛，可以捕获外界光刺激并转换为红外光刺激，为NR600假体提供电能和传输信号。NR600的眼内植入部分尺寸为3×4mm，包括了600个电极，电极长度为150μm、间距为100μm。目前已在9名患者中进行植入，能够恢复病人对光刺激的感知，此外有一名病人的光栅视锐度测试结果≤ 0.1cpd，对应的视力为20/6000。目前NR600视网膜下假体的临床试验还在进行中[167]。

近年视网膜下植入物也已进入临床试验，即斯坦福大学开发的PRIMA视网膜下假体。该植入物使用专用护目镜，投射近红外光来激活和驱动能够直接刺激双极细胞的硅基底的光电二极管，即可以通过无线的方式，直接通过光线将能量和视觉信息传递到视网膜下的像素阵列。像素阵列独立的植入于视网膜下，不需要在巩膜等位置建立永久性的通道。相较于传统的巩膜切口，避免了导电装置性能故障和慢性炎症反应，同时也避免了跨眼电路，降低了手术难度。临床研究使用的像素阵列尺寸为2×2mm，包括378个像素，每个像素的直径为100μm。2022年，法国公司 Pixium Vision获得相关技术授权并继续开发，将其推进到临床试验阶段[168]。2024年，Science Corporation宣布收购了PRIMA产品，并于2025年6月向欧盟公告机构提交PRIMA视网膜植入系统的CE标志认证申请。PRIMA[169]已经在欧洲完成了试验PRIMAvera （NCT04676854），共入组38 名患有地图样萎缩（GA）的受试者。所有参与患者的视力在植入后12个月通过logMAR评分进行测量。结果显示：与基线相比，植入后12个月平均改善23个字母（4.6行），提高最多的患者最大改善为59个字母（11.8行）[170]。

前期研究中，有5名受试者进行了PRIMA视网膜下假体植入手术，其中3名患者植入位置位于黄斑区域的视网膜下，该三名患者的视力为20/460至20/550，另外有一位在黄斑周围植入假体的患者视力为20/800[171-172]。目前植入PRIMA视网膜下假体的患者最佳视力为20/460，覆盖的视野范围约为17 deg。通过降低像素尺寸，使用每个像素直径为70或50μm的像素阵列，理论上可进一步提高视力。PRIMA视网膜下假体恢复的视野范围主要受到两方面限制，其一为刚性的基底材质，不允许高曲率的弯曲，难以贴合视网膜的球形形状。此外，即使多个刚性单元可以覆盖一个较大的弯曲区域，巨大的视网膜下植入阵列会增加视网膜网脱的风险。

国内复旦大学张嘉漪等人在2018年已开发出金纳米颗粒装饰的二氧化钛纳米线阵列（Au-TiO2 NW阵列）作为人工光感受器，并完成小鼠、猴子等动物试验[173]，于2023年开始临床研究。2025年，上海多家科研团队合作开发出全球首款覆盖可见光至近红外二区（470-1550nm）的超宽光谱视觉假体，不仅让因视网膜退化而失明的小鼠重获可见光视觉，更赋予其感知红外图案、识别红外信号的能力[174]。

（3）脉络膜上假体（Suprachoroidal Prosth）

这类假体放置在脉络膜与巩膜之间，避免了对玻璃体腔的干扰，机械固定更为稳定，且散热性能相对良好。不过，由于电极与双极细胞之间仍有一定距离，所需刺激电流更大，有潜在损伤组织的风险，实际视觉分辨率也相对有限。澳大利亚和日本的研究团队均已开展相应的临床试验。

澳大利亚眼科研究中心开发了一种脉络膜上假体（BVA），该假体具有44通道，以交错网络排列在19×8mm的硅胶基板上，该电极阵列能够覆盖10×7.5mm的视网膜区域，对应于约37.6×27.6deg的视野范围。相较于第一代，第二代脉络膜上假体覆盖了更大的视网膜范围，同时提高了电极阵列的密度，以期望获得更高的分辨率。在2021年刚刚公布了中期临床试验结果。在该临床试验中，共有4名晚期的RP患者进行了假体植入。在植入电极阵列56周后，98%的电极可以正常工作，电极-视网膜组织界面保持长期稳定且阻抗稳定，未发生严重的不良事件。4名患肢均能够稳定的诱发光幻视，2名患者在使用假体时，对运动物体的识别和空间分分辨率测试实验中表现显著好于未使用设备时的表现。4名患者均在使用假体时对触摸桌面物体、障碍物检测等测试中表现更准确。尽管如此，4名患者的视力恢复水平仅停留在有光感的水平，空间分辨率等仍需进一步提升决。假体在3名患者眼内12个月内均保持稳定，没有发生横向移动，同时在3名患者中均能够可靠的引发光幻视。此外在光定位测试中，3名患者均能以显著高于随机几率的正确率识别刺激。此外，3名患者中仅患者1能够完成视锐度测试，该名患者的恢复视力虽高于未使用假体装置的视力，但其恢复的效果仍然很差，需进一步提升，如提高电极数量与密度等[175]。日本Fujikado等人在两名晚期的RP患者的各一只眼内植入脉络膜上假体（Suprachoroidal-Transretinal Stimulation, STS），系统包括电极阵列和内部装置，电极阵列部分的尺寸为5.7×4.6mm，共49根直径为0.5mm、间距为0.7mm的电极组成，电极阵列部分植入到眼睛的巩膜与脉络膜之间，系统的部装置植入头部颞侧的皮下。在第一代产品中，有源电极数量仅9个，在第二代产品中增加到了49个。在植入第一代STS后4周，通过刺激特定电极，能够引起病人的光幻视，在搭配摄像机作为视觉对象的检测器时，患者可以在黑色背景下识别和找到白色目标刺激。植入的STS假体系统能够在4周内保持安全性和稳定性，但是由于需要跨眼电路等，在植入STS后，两名患者在各个方向的眼球运动均受到轻微限制[176]。在3名RP晚期患者植入第二代STS后1年时间，所有RP患者均能够稳定诱发光幻视，电极阵列植入靠近中央凹的患者执行视觉任务的改善更大[177]。

3.2 皮层视觉脑机接口

皮层视觉脑机接口（Cortical Visual Prosthetics）是一种直接作用于大脑视觉皮层的神经接口技术，旨在通过电刺激皮层神经元以重建盲人的视觉感知。该技术跳过了眼球和视神经等前端结构，适用于视网膜和视神经严重损伤但视觉皮层功能相对保留的患者，尤其对先天性失明不适用而皮层仍具可塑性的人群提供了理论上的治疗可能。

早期皮层视觉脑机接口研究主要采用皮层表面的脑电图电极（ECoG electrodes）进行刺激，这种电极具有相对较低的侵入性和较大的电极接触面积，可用于初步探索皮层刺激所产生的感知特征。但由于ECoG电极在空间分辨率和刺激聚焦度上的局限，同时有诱发癫痫的风险，后续研究团队转而使用硬质微针型电极阵列（如Utah array）进行皮层内刺激，希望获得更高密度、更精细的感知调控。进入21世纪后，随着材料科学的进步，超柔性电极的出现为视觉皮层刺激带来了新的可能。这类电极具备极高的柔韧性和良好的生物相容性，能够紧贴脑组织表面或实现多通道皮层内植入，减少组织损伤与异物反应，并显著提升刺激的安全性和稳定性。

皮层视觉脑机接口的刺激目标通常包括视觉皮层的各级脑区，因这些区域在视觉空间的拓扑映射上具有高度规则性。通过系统建立图像与刺激模式之间的对应关系，研究者试图以点状光感（phosphenes）的形式重建空间视觉图样。目前已有多项研究在动物模型和人类患者中验证了皮层刺激诱发光感的可行性，并逐步探索更高阶感知维度的编码方式，如颜色、运动和形状等。但该类技术因其侵入性特征，在手术安全性、电极寿命、刺激特异性等方面仍面临诸多挑战，尚处于临床前或早期试验阶段。未来，随着柔性高通道电极、无线传输系统以及人工智能辅助解码技术的发展，皮层视觉脑机接口有望实现从“点感知”到“图像理解”的跨越，成为高阶视觉功能重建的核心路径之一。

（1）Orion。由Second Sight Medical Products, Inc.开发的Orion视觉皮层假体，由摄像头、计算机和60个表面电极组成的硬膜下阵列组成。2017年底开始的临床试验包括五名盲人患者，计划进行为期5年的随访[178]，已获初步结果[179]。

（2）ICPV项目。伊利诺伊理工学院普利兹克生物医学科学与工程研究所执行主任、生物医学工程教授Philip R. Troyk领导、多机构参与的皮层内视觉脑机接口（Intracortical Visual Prosthesis, ICVP）项目使用无线浮动微电极阵列（Wireless Floating Microelectrode Array, WFMA），阵列由放置在视觉皮层表面的16个聚对二甲苯绝缘铱微电极组成，还包括集成电路微处理器、无线电源和功能激活的微线圈。项目开发的Intracortical Visual Prosthesis System中，安装在眼镜或头带上的摄像机连接到视频处理器单元，该单元将图像转换为映射到电极阵列的图案。信号通过刺激模块传输到位于头部的遥测控制器。该系统正在进行人体临床试验[180]。在临床前阶段，伊利诺伊理工大学团队与拉什大学医学中心的神经外科医生合作开发和改进外科手术，最终成功地将25个刺激器和400个电极植入失明患者体内。临床阶段旨在测试该假体是否将为研究参与者提供更好的导航和执行基本的视觉引导任务的能力，将在芝加哥开展临床试验[181]。

（3）Gennaris。由澳大利亚墨尔本莫纳什大学视觉团队开发的Gennaris包括安装在眼镜上的摄像头，用于捕捉场景并将其传输到“袖珍处理器”，处理器提取有用的信息，然后将其发送到植入大脑初级视觉皮层第4层的电极中。信号由位于头部后部的无线发射器广播[182]。对实验动物的安全性测试证实其有效果，组织学检查显示植入后对皮层的损害最小，并且可进行长期刺激而不会发生不良事件[183-184] ，计划在未来几年内进行首次人体临床试验[185]。

（4）Blindsight。Neuralink的Blindsight项目是基于高密度柔性微电极阵列的视觉脑机接口系统，旨在通过在视觉皮层直接植入并刺激微电极，绕过受损视网膜和视神经，实现盲人视觉重建。Blindsight 已在猕猴模型中成功应用，其核心硬件包括单根线最大32通道、单套系统最多可植入96根线、总通道数高达 3072通道的高密度电极阵列，并配合机器人辅助的微创高精度植入技术。动物实验表明，通过视觉皮层的多点同步电刺激，能够稳定诱发目标位置的伪视觉表征，使实验动物在约2/3试验中根据电刺激“看见”并定位虚拟目标。在临床应用方面，2024年 Blindsight已获得FDA“突破性医疗器械”认定，预计2025年底将率先在阿联酋和美国启动首批植入试验，主要面向重度视觉障碍患者。临床方案包括通过3072通道高密度柔性电极阵列实现初级低分辨率（类似于“像素化”的Atari级图像）视觉输出，并逐步提升算法和硬件密度，向更高分辨率和更丰富的视觉感知目标迈进。该项目虽然在技术可行性、长期稳定性、伦理安全性等方面仍面临挑战，但仍有望成为全球首个大规模推进盲人视觉重建的视觉脑机接口临床应用系统[186]。

此外，还有一些团队在开发皮层视觉脑机接口，并获得初步结果。西班牙埃尔切米格尔·埃尔南德斯大学神经工程主任Eduardo Fernandez团队开发的皮层视觉脑机接口（电极），植入盲人患者体内，让患者获得基本视力，但其植入时间最多6个月，因为假体还未被批准长期使用[187]。还需要进一步验证该假体的电极能持续多久发挥功能，因植入大脑后人体免疫系统会分解电极并用疤痕组织包围它们，最终削弱信号。研究人员预计该脑机接口/假体可能会持续两到三年，甚至可以持续10年。该假体就像人工耳蜗一样，需要通过颅骨无线传输其信号和电源以到达电极。美国贝勒医学院Daniel Yoshor教授团队开发的皮层视觉假体通过动态电流电极刺激大脑皮层，成功在受试者脑海中呈现指定的图像[188]。研究人员对传统的电极进行了改进，结合了电流转向和动态刺激，通过精准的电流控制，依次激活不同的电极，来实现字母或图片轮廓的绘制。该仪器在6名受试者中验证，结果表明，6 名受试者都“看到”了字母“Z”的存在，盲人的确可以通过使用将视觉信息直接输入大脑的技术，来重新获得识别视觉信息的能力。下一步，研究团队将和神经工程师合作，开发具有数千个电极的电极整列，改进刺激算法，从而能够精准地刺激大脑视觉皮层的每一个神经，帮助盲人“看到”更复杂的信息。荷兰神经科学研究所（NIN）Pieter Roelfsema团队开发了一个高分辨率的神经假体，包含1024个以类似矩阵模式排列的皮层内电极。将这个电极装置植入了两只猴子的大脑视觉皮层，使受试猴子大脑中产生可识别的运动和复杂形状的感知，包括朝向、字母形状和运动感知[189]。犹他电极的发明人Richard A. Normann和Pieter Roelfsema团队合作，在人的视觉皮层植入了96通道的犹他电极阵列，在6个月的短期临床研究中，实现了英文字幕的写入和黑白边界的识别[190]。

国内有多个研究团队致力于或已经启动皮层视觉脑机接口的研究：临港实验室分三阶段，规划了面向临床1280，5120，20000通道的新型皮层视觉脑机接口器件的研发。目前已经在非人灵长类猕猴上植入了高密度神经电极，实现了基本视觉形状在皮层的写入，电刺激诱导猕猴开车导航到目标位置等目标任务。北京大学方方团队在人上使用SEEG电极开展了一些电刺激诱发光幻视的研究。

- 脑疾病调控

脑功能性疾病已成为全球主要的致残和致死原因之一，不仅给个人、家庭和社会带来了沉重的负担，其治疗和机制研究也构成了巨大的科学挑战。在脑功能性疾病与精神疾病给全球带来严峻负担的背景下，借助脑机接口和神经调控技术的快速发展，脑机交互神经调控已成为国际学术和应用研究的前沿和热点方向。技术进步为脑健康干预、脑疾病治疗以及脑功能康复提供了突破性的解决方案，展示了巨大的临床潜力和应用前景。

4.1 运动障碍

深脑刺激（Deep Brain Stimulation, DBS）是一种通过植入大脑的电极向靶点核团施加电刺激并调控脑功能的神经外科疗法，在帕金森、特发性震颤、肌张力障碍等运动障碍的治疗中已应用超30年。Benabid研究团队在1987年首次将DBS用于运动障碍的治疗以来，该技术在临床上得到了广泛应用，全球植入量超过244,000例[191]。帕金森病是DBS应用最成熟的适应症之一，以STN或苍白球内侧部（Globus Pallidus interior, GPi）为靶点的高频刺激可以有效改善患者的震颤、僵直和运动迟缓等症状。一项随机对照研究表明，与单独药物治疗相比，DBS具有更显著的运动症状改善和生活质量的提高[192]。在特发性震颤的治疗上，丘脑腹中间核（Ventral Intermediate Thalamic Nucleus, VIM）已成为DBS最常用的靶点。VIM-DBS对药物治疗无效的特发性震颤患者具有良好的治疗效果，可明显减少上肢、头部和声音震颤，提高患者的日常生活能力[193]。DBS 还被应用于肌张力障碍，特别是原发性肌张力障碍患者，GPi和STN是DBS治疗肌张力障碍的两大主要靶点。研究表明，GPi-DBS和STN-DBS在改善运动症状、生活质量具有相似的疗效，但STN-DBS可能会引发更高的刺激相关不良反应，如情绪波动和抑郁[194]。

在现有的疗法体系上，临床上也不断探索DBS的新型刺激模式和刺激靶点。新型刺激疗法如变频刺激相对于传统的高频刺激对帕金森病患者的步态等中轴症状具有更好的改善效果[195]。与此同时，在经典的VIM、STN和GPi等经典靶点外，临床上逐渐探索了大脑脚间核（Pedunculopontine nucleus, PPN），苍白球外侧部（Globus Pallidus external segment, GPe）等新的刺激靶点，以进一步改善DBS在步态平衡、睡眠障碍等方面的疗效[196]。对于特发性震颤，一些临床实验结果表明后丘脑底区后部（Posterior subthalamic area, PSA）和未定带（ZI）区域可能比VIM更有效地抑制震颤。

DBS设备的工程技术发展进一步提高了DBS治疗运动障碍的疗效。在可充电技术和远程程控技术逐渐成熟并广泛应用的同时，近年来DBS设备技术在方向电极、闭环刺激方面取得重要进展。方向性电极的应用更精准地靶向病理区域并减少非目标区域刺激引起的副作用[197]。闭环刺激相对于传统的开环刺激能实时根据患者病症和生理状态调整刺激参数，从而具有更好的症状控制效果和更高的刺激效率等优势[198]。而上述新靶点、新刺激模式的拓展离不开对于疾病机制以及调控机制的深入理解。因此，进一步深入机制探索至关重要。

4.2 难治性癫痫

作为高发病率的神经退行性疾病，难治性癫痫患者因药物耐药性需寻求外科干预，但传统外科治疗的适用人群受限。癫痫是一类由于大量神经元异常的同步化放电所致的常见神经功能性疾病，我国的患病率约7‰[199]。癫痫的临床治疗目前虽仍以药物治疗（包括生酮饮食）为主，但尚有30%～40%的癫痫患者无法从中获益，需要外科手术干预，这类患者被称为“难治性癫痫”[200-201]。而针对难治性癫痫的外科治疗手段，大致可分为四类：传统的开颅切除性手术；传统的开颅姑息性手术；微创毁损性手术；神经调控。其中，传统手术（主要包括癫痫病灶的切除或癫痫网络的离断）和微创毁损性手术虽然对于一部分患者效果良好，但存在相对较大的神经功能损伤风险，不少患者术后会遗留长期且明显的记忆、认知等高级神经功能障碍[202-203]；而对于部分癫痫灶广泛，或癫痫灶涉及肢体运动、语言等大脑显性功能区等患者，传统手术和微创损毁性手术是无能为力的，此时就需要考虑神经调控这一新型的治疗方式。

基于闭环电刺激的RNS与靶向丘脑的DBS已逐步取代传统高风险手术，为患者提供微创治疗新方式。目前，临床针对难治性癫痫的神经调控方式主要包括VNS、DBS与反应性电刺激（Responsive Neurostimulation, RNS）等。VNS是一种非侵入式治疗方式，也是诸多神经调控技术中相对效果较弱的一种。癫痫改善率，即术后癫痫发作频率或发作强度降低一半以上的比率，约为50%。VNS通过将刺激电极触点缠绕在颈部迷走神经上来进行神经调控，临床主要应用于儿童全面性难治性癫痫[204-206]；而DBS与RNS均属于侵入式神经调控技术，也是近年来及未来难治性癫痫的主要神经调控方式。DBS在难治性癫痫中的治疗主要基于癫痫网络的“传播点”理论，该理论认为，部分癫痫发作时大脑异常神经元放电会经由某一个或几个传播点后投散至大脑皮层，引起较明显的临床症状。因此通过在传播点永久植入脑深部电极并采取电刺激的方式，可有效调控癫痫异常放电特性，从而起到治疗作用。目前在难治性癫痫中常用的DBS靶点包括丘脑前核（Anterior nucleus of thalamus, ANT）、海马（Hipp）、丘脑中央核（Centromedian thalamic nucleus, CMT）和STN等[207-208]。DBS治疗难治性癫痫的效果往往在术后数月内即可达到稳定的高峰值，配合良好的参数程控设置，DBS治疗的癫痫改善率约56%~75%。相比于VNS和DBS这一类开环性质的神经调控，RNS是基于闭环刺激的神经调控新技术，2023年美国FDA批准由美国NeuroPace公司研发的RNS®系统用于治疗成人的局灶性难治性癫痫，而我国自主研发的首款RNS神经调控设备EpilcureTM也于2021年顺利实现首例人体植入并随后进入正式临床试验。RNS技术主要是将电极分为两种属性：其一是脑电信号采集，将电极置于癫痫发作起始区域并予以24小时实时信号采集与检测，对可能存在的癫痫样异常脑电信号进行预警；其二是电刺激治疗，将电极置于癫痫发作起始区域或传播点区域，类似DBS方式，通过电刺激来调控癫痫异常放电，起到治疗作用。不同于DBS的脑深部电极模式，目前RNS技术采用的电极包括脑皮层电极和脑深部电极两类。由于采用了闭环调控模式，RNS技术颠覆了难治性癫痫的经典神经调控理念，可以实现类似“反导系统”样的精准刺激和智能调参，其治疗效果也呈现逐渐增强的特点，现阶段临床治疗结果显示，RNS治疗前三年的癫痫改善率在50%~60%，而在治疗第7~9年的癫痫改善率在70%以上，整体癫痫改善率在53%~75%；且相比于DBS，RNS治疗的能量消耗更低、副作用更小[209-211]。

4.3 难治性抑郁症

抑郁症作为全球公共健康重大挑战，难治性患者无法从传统的药物、心理、物理治疗中获益，其高致残率与自杀关联性凸显了治疗手段创新的迫切性。随着社会应激因素增加及人口老龄化，抑郁症发病率逐年攀升，无论在全球还是在中国，抑郁障碍在精神障碍中的疾病负担均居首位[212]，我国成人抑郁障碍终生患病率高达6.8%[213]，其中约30%的患者会因药物、心理等传统治疗方式无效或者效果欠佳而成为难治性抑郁（Treatment-Resistant Depression, TRD）[214]。每年约28万的自杀人口中40%患有抑郁症，凸显了对有效治疗手段的迫切需求。难治性抑郁症作为一种严重的精神健康问题，对个体的生活质量、工作效率以及社会功能造成深远影响，其治疗手段的创新与进步直接关系到国民健康水平的提升和人力资源的优化配置。美国阿拉巴马大学伯明翰分校（UAB）在难治性抑郁症治疗领域的最新进展是启动了名为TRANSCEND的脑深部电刺激（DBS）临床试验。该研究针对的是至少对四种抗抑郁药无效的患者，通过向大脑中与情绪调节相关的胼胝体下扣带回发送温和电脉冲来进行干预。试验采用严格的12个月随机双盲、假刺激对照设计，以科学验证其疗效和安全性。所采用的DBS设备已获得美国FDA的突破性设备认定，有望为最严重的难治性抑郁症患者提供一种全新的治疗方案。上海交通大学医学院附属瑞金医院联合复旦大学类脑智能科学与技术研究院、剑桥大学等机构，利用颅内电生理记录结合机器学习，发现了能够预测治疗效果的客观神经生理标志物，开展了针对26名难治性抑郁症的深部脑刺激（DBS）临床试验，50%的患者达到了治疗响应标准（HAMD评分下降≥50%），35%的患者实现了临床缓解（HAMD评分<8），为实现抑郁症的精准化、个性化治疗提供了强有力的科学依据。杭州青石永隽医疗设备有限公司以难治性重症抑郁为切入口，突破“抑郁情绪标记物识别”，“最优植入靶点选择”，“自适应闭环调控”等临床技术难点，开发自主知识产权的脑机接口诊疗系统，目前一代产品已完成开发，通过内测动物实验，进入型检阶段，诊疗方法已完成40余例急性临床验证和4例长期科研型临床，4例长期科研临床100%有效，75%治愈。

4.4 难治性强迫症

强迫症作为致残率超50%的精神疾病，传统疗法存在治疗耐药性强、副作用显著及资源可及性差三大困境。强迫症是一种慢性、致残性精神障碍，其特征为反复侵入性思维（强迫思维）及相应的强迫行为，给患者带来显著的焦虑与痛苦[215]。该疾病通常起病于儿童、青少年或成年早期，终生患病率约为2.4%[216]。强迫症不仅消耗患者大量精力，严重影响生理健康、社交功能及生活质量，同时显著增加自杀风险——5%~15%的患者曾尝试自杀，1%~2%因强迫症相关因素死亡[217]。国内报道，其卫生经济支出每年约370亿元，人均直接和间接经济支出2-3万元[218]。强迫症号称精神疾病的癌症，通常表现为慢性病程，病程长达数十年，给患者个人家庭及社会带来了严重的精神和经济负担。现阶段，精神药物治疗及认知行为疗法（Cognitive Behavioral Therapy, CBT）是国际国内强迫症治疗指南推荐的一线治疗方法[219-220]。然而，强迫症的治疗面临诸多挑战。CBT依赖专业治疗师，且受限于可及性与经济负担，尤其在资源有限地区更为突出。药物治疗的副作用也较为明显，包括头痛、恶心、失眠、性功能障碍，甚至可能导致心律失常、抑郁加重、戒断反应和长期代谢问题。部分患者即便接受充足剂量和疗程的药物治疗和/或CBT治疗，疗效仍不理想，约20%的患者发展为难治性强迫症[221]，亟需寻求更有效的干预手段。

DBS逐渐应用于难治性强迫症的治疗[222]。截至目前，全球已有超过300名强迫症患者接受了DBS植入手术[223]。多项研究表明DBS治疗难治性强迫症疗效确切且远期疗效稳定。DBS治疗的安全性已经在运动障碍中得到了充分的验证，而强迫症患者多为中青年，手术耐受性更佳。目前，DBS治疗强迫症已得到国内外监管机构和学会的广泛认可。2009年，FDA在人道主义器械豁免（Humanitarian Device Exemption, HDE）框架下批准DBS以内囊前肢（Anterior Limb of Internal Capsule,  ALIC）为刺激靶点，用于治疗强迫症[224-225]。同年，DBS治疗强迫症在欧洲获得CE认证。在2014年，世界立体定向与功能学会召集国际专家组发布了精神疾病立体定向神经外科共识指南[226]。该指南采用了精神科医生同意的难治性强迫症的定义作为DBS的适应证[227]。2022年，世界生物精神病学协会联合会指出，DBS可应用于慢性、严重且对传统治疗无效的强迫症患者。2024年，中华医学会精神病学分会精神外科联盟发布了《难治性强迫症脑深部电刺激疗法中国专家共识》[228]，进一步推动了DBS技术在强迫症治疗中的规范化应用。这一系列国际和国内的共识确立了DBS在强迫症治疗中的临床价值。相较于难治性强迫症对患者所造成的严重危害，DBS手术对强迫症患者来说是一项“拯救性”手术。与其他神经外科手术相比，DBS手术的风险显著较低且可接受。DBS手术潜在的获益明显超过其可能的风险[229]。

# （二）非侵入式脑机接口临床应用

- 检测与交互脑机接口

非侵入脑状态检测、交互系统的研究取得了显著进展，其在以意识、睡眠、情绪和注意力为代表的多个领域展现出广阔的应用前景，为医疗、心理、人机交互等领域提供强大支持。

1.1 意识评估脑机接口

基于脑电EEG的脑机接口技术能够不依赖外在行为表现直接检测神经行为响应，易于重复采集，可床边操作，经济成本较低等优点，因此被用于意识障碍（Disorders Of Consciousness, DOC）患者的诊断、康复预测与神经反馈治疗[230]。脑机接口可提供丰富的任务态范式，实时获取与任务相关的更加稳定的脑活动模式的改变，通过实时的神经反馈增强具有鉴别意义脑信号特征，准确评估意识水平，有效地识别认知运动分离的患者，从而进行有效干预并促进意识康复[231-233]。任务态的脑机接口可以获取意识障碍患者在执行任务时的大脑神经活动信号，直接检测患者中与任务相关的EEG信号模式的改变并作为识别认知运动分离患者的依据，为诊断与预后评估提供可能的方法。如EEG-BCI评估无反应综合征患者的命令跟随能力，若患者在任务中能够按照指令完成想象“打网球”或“在房间内走动”的动作，并产生与健康被试无显著差异的神经行为响应，则表明该患者虽然临床诊断为无行为反应，但能够理解指令并完成想象，因此被认为是具有命令跟随能力的认知运动分离患者，由此减少临床行为量表的误诊[234-235]。EEG-BCI的时间分辨率高，能够捕获事件相关诱发电位ERP，可用于特定感觉通路的完整性评估，为弥漫性脑损伤的整体影响提供可靠的信息[236]。此外，EEG-BCI可以实时记录大脑对扰动的响应信号，可与干预方法形成闭环，从而指导治疗[237]。比如对患者进行直流电刺激或经颅磁刺激时，可根据EEG响应实时调整参数。

浙江大学研究团队在2018年提出了在特定听觉刺激状态下的定量的脑电图指数（Quantitative EEG, QEEG），可以作为意识障碍患者的预后指标[238]。2019年，美国哥伦比亚大学团队研究检测意识障碍患者的EEG信号中对命令（比如移动手）响应的脑活动。结果显示，在104位急性脑损伤导致的无反应患者中，15%的患者对命令缺乏行为反应，但在其EEG信号中可以检测到对命令响应的脑活动模式。而这类CMD患者的预后结果更好[239]。2020年，中科院脑科学研究所采用基于语言范式的EEG信号追踪语音层级加工过程，并可靠地预测单个患者的康复结果[240]。意大利巴勒莫大学团队采用触觉脑机接口研究认知过程对患者评估的作用，帮助医生预测意识障碍患者的康复结果[241]。2020年，华南理工大学研究团队采用多模态的脑机接口有效地识别出认知运动分离的患者，并发现这类患者的预后结果显著优于其他患者[242]。2022年，g.tek研究团队[243]研究表明脑机接口在意识障碍中的应用主要分为评估、交流、预测和治疗四个方向，脑机接口预测康复结果对于急性和亚急性期的患者至关重要，因为预后直接影响治疗方法的选择，甚至影响医生和家属对维持生命的决策。团队还总结了多年的研究并认为脑机接口技术作为一种非侵入式的神经反馈方法，在意识障碍的治疗中展现出了广阔的应用前景[244]。2024年，首都医科大学神经外科团队从临床角度再次证实脑机接口能为意识障碍患者的意识与功能的恢复提供新的途径[249]。

目前用于意识障碍的脑机接口大多基于某个特定任务诱发的脑电信号，直接诊断患者的意识水平或进行预后评估[245-247]。然而，这种方法存在一定的局限性，主要表现在对患者的认知负荷要求过高、对多个感觉通路完整性的依赖，以及对持续注意能力的要求，导致评估结果中存在较高的假阴性率。此外，由于缺乏神经反应机制的深入理解，评估结果的可解释性仍显不足。

1.2 睡眠脑电监测

睡眠的监测和评估是睡眠医学和临床研究的关键组成部分，也是预防、诊断和治疗睡眠障碍的基础。睡眠过程中的EEG信号能够直接反映不同睡眠阶段的大脑活动状态，通过监测睡眠脑电可以直接且高效地进行自动睡眠分期以及定量计算各阶段时长等信息，从而准确、客观地评价睡眠质量。

作为睡眠监测的重要应用方向，基于脑电信号的自动睡眠分期技术近年来实现了令人瞩目的进展。该技术借助机器学习、深度学习算法的引入，大幅提升了数据分析效率，有效减轻了人工标注过程中主观因素的干扰[248]。纽约大学团队提出利用多尺度卷积神经网络提取时序及频谱特征，并通过双向长短时记忆网络（BiLSTM）捕捉睡眠阶段转换规律，从而取得了较高的分类准确率[249]。帝国理工大学设计的XSleepNet模型采用多模态信号，分别利用CNN和BiLSTM网络从一维EEG信号及二维时频图像中提取特征，再经由BiLSTM/BiGRU层进行序列信息处理[250]。浙江大学团队提出了一种基于卷积和注意力机制的神经网络，使用结合注意力机制的卷积神经网络作为特征提取器，并用多个注意力层替代循环神经网络架构，在单通道EEG上取得了不错的自动睡眠分期效果[251]。北京交通大学提出了动态图卷积神经网络可以更好地提取空间特征来捕捉不同睡眠阶段之间的转换规则[252]。北京理工大学团队提出了基于BiLSTM的轻量级自动分期模型，在较小的参数量条件下取得了相似的性能效果[253]。

除了睡眠自动分期，一些研究者聚焦于不同的睡眠阶段具有的微小脑部特征，也就是监测睡眠中出现的特征波。瑞士弗赖堡大学设计了睡眠纺锤波检测算法MUSSDET，通过计算一段EEG的可测量特征来客观地描述脑电信号的轨迹，并使用SVM进行分类来识别睡眠纺锤波[254]。武汉大学团队提出了一种基于深度神经网络的框架，通过混合深度特征（微尺度）和熵特征（宏观尺度）来自动检测睡眠纺锤体[255]。北京师范大学团队提出特征融合网络与脉冲神经网络（SNN），结合深度学习与传统信号处理技术，优化睡眠纺锤波的检测精度[256]。复旦大学团队提出了一种基于单通道EEG的两阶段睡眠纺锤波检测方法。使用自适应参数的Teager能量算子发现稳定数量的睡眠纺锤波候选者，再设计代表性特征并利用bagging分类器从所有候选者中进一步识别真实纺锤波[257]。天津大学团队提出类似的分层纺锤波检测算法，第一层使用Morlet小波和RMS方法检测纺锤波，第二层采用改进的k-means算法提高纺锤波检测效率[258]。

近年来，无创神经调控方法已经成为医生和睡眠障碍、抑郁症等相关疾病患者的主要选择之一。意大利卡塔尼亚大学发现重复性经颅磁刺激（rTMS）有改善睡眠质量方面的作用[259]。美国麻省理工学院提出经颅交流电刺激（tACS）通过节律同步机制能够调节大脑的神经振荡，进而影响睡眠的结构和功能[260]。电子科技大学的研究也表明低频tACS可能触发睡眠稳态的自上而下处理过程，加速睡眠的起始[261]。南京大学团队提出tACS可与深度学习、脑机接口等技术结合，实现脑电节律的实时干预和个性化治疗，并考虑在航天员睡眠失调中应用[262]。

1.3 情感脑机接口

情绪是人对客观事物和主观认知在大脑中形成的一种心理状态，通过检测大脑信号可以观察和评估人的情绪。通过脑机接口系统检测情绪相关的脑信号是当前研究情绪的主要手段，现有的情绪脑机接口有使用脑电图EEG或功能近红外光谱成像fNIRS的单一模态系统，也有将多种脑信号融合的多模态系统[263-264]。

深度学习方法已被广泛应用于情绪脑状态解码的研究。电子科技大学脑器交互团队提出一种基于脑电网络认知先验图的情绪脑电深度图学习系统（BF-GCN），能有效挖掘情绪脑电信号的潜在图模式，实现稳健的情绪脑电解码[265]。迁移学习思想有助于解决EEG情绪识别任务中的个体差异和数据稀缺问题，清华大学团队针对便携式家用级脑电设备，通过知识迁移提升从少通道EEG信号中进行情绪识别的性能[266]。天津大学团队针对跨数据集的脑信号解码问题，提出了结合流行嵌入知识迁移与自适应批归一化方法的跨域适应技术，提升了模型跨数据集的迁移性能[267]。意识障碍患者的脑状态解码对其诊断治疗有重要意义，华南理工大学与琶洲实验室团队提出一种自监督对比域泛化框架SSCDG，该模型可有效地解码由音乐诱发的意识障碍患者的情绪相关脑状态[268]。大模型在情绪识别和脑状态解码领域展现出重要潜力，跨模态预训练模型通过融合多模态数据提升了情绪表征和检测的能力。伊利诺伊大学与哈佛医学院团队通过大规模、多模态的脑电信号预训练出BIOT模型，具备对情绪、睡眠等多模态脑信号的检测能力[269]。

韩国首尔三星医疗中心团队运用可穿戴设备开展抑郁症相关的生理指标和行为模式的持续监测，该技术为及时发现抑郁症提供可能，并提供及时的干预措施[270]。中国科学院自动化所研究团队研发了一种可穿戴rTMS设备，可与脑电监测系统集成，实现实时反馈与自适应调节，从而支持闭环神经调控，提升个性化治疗效果[271]。台湾地区国立交通大学团队针对智能家居交互场景，提出用单通道EEG实时监测用户认知状态（如警觉/困倦），并联动家庭网络自动调控灯光等设备，为用户提供个性化服务[272]。针对晕动症，华南理工大学与琶洲实验室团队开发了一种正念脑机接口接口系统，在31名易患晕动症的受试者参与的真实汽车行驶环境实验中，所提出的正念脑机接口能够有效缓解晕动症症状，是首个被证实对晕动症有缓解效果的非药物方法[273]。重庆大学与密歇根大学团队研究了驾驶对不同情绪下大脑功能连接网络的影响，发现在参与驾驶后，不同情绪状态下大脑区域间功能连接的强度显著增强，功能连接网络中的信息传递更加高效、快速[274]。情绪驱动的沉浸式体验与虚拟角色互动吸引了娱乐游戏领域开发者的关注。德国普朗克研究所团队运用沉浸式虚拟现实（VR）技术进行自然刺激研究情绪唤醒的脑信号特征，发现了情绪唤醒与顶枕叶α波功率之间的关联[275]。

1.4 注意力脑机接口

注意力是人类大脑的高级认知功能，在日常生活中发挥着重要作用，使个体能够有效地分配认知资源并专注于特定任务和刺激。基于EEG的注意力检测与识别已成为BCI研究中的重要方向之一，受到广泛关注。早期研究多采用特征工程方法，即从EEG信号中提取预定义的时域、频域和非线性动力学等特征，并结合传统机器学习分类器进行识别[276-277]。加利福尼亚大学团队于1997年提出了基于EEG功率谱变化来表征警觉和注意力水平，是最早使用频域特征进行注意力识别的工作之一。台湾交通大学团队为了持续评估驾驶员注意力的集中程度，分别计算各个脑区的功率谱特征并使用独立成分分析的方法进行特征选择，最后使用SVM取得了86.2±5.4%的二分类准确率[278]。兰州大学团队融合了EEG信号时域、频域和非线性特征，并提出了基于相关性的特征选择方法，对注意力高、中、低三类状态的识别准确率达到了80.84±3.0%[279]。

近年来为了克服EEG的个体差异和增加BCI系统实用性，基于多模态、大数据和深度学习方法的跨个体注意力解码取得了显著进展[280]。随着电极材料、可穿戴设备以及大模型等技术的发展，基于EEG的注意力检测识别逐步拓展到注意力调控和辅助临床康复[281-282]。2024年，华南理工大学团队提出了一种端到端的跨个体注意力解码方法DA-TSent，使用域对抗训练以及时空一维卷积神经网络，减少个体差异以及注意力状态与非注意力状态嵌入特征之间的重叠，实现了高效的注意力在线识别，结果显示其在线注意力识别准确率可达89.02%±9.58%，高于现有基线方法。团队还提出了一种用于实时注意力调节和认知监测的可穿戴BCI系统[283]，结果表示可穿戴BCI系统能够实现注意力的实时、准确监测，长期神经反馈训练能够显著提高受试者注意力调节能力。上海大学团队提出了一种基于EEG和眼动的多模态生理信号的检测视觉注意力方法，并应用于轻度认知障碍（MCI）患者的自动、非侵入性且快速高效的筛查[284]。在152名MCI患者和184名健康老年对照组上的实验结果显示，筛查准确率达到81.51%，能够成为对大规模人群进行个性化精准MCI筛查的有力工具。此外，基于注意力BCI系统的神经反馈训练对注意力缺陷/多动障碍（Attention Deficit Hyperactivity Disorder, ADHD）患者的干预和行为认知治疗具有重要价值[285-287]。德国波茨坦大学团队设计了随机的单盲对照试验，验证基于EEG的神经反馈对减轻ADHD患者症状的有效性。共有35名ADHD儿童随机分配到神经反馈治疗组和对照组，实验结果表明基于EEG的神经反馈训练能够显著减轻ADHD症状，对患者注意力提升和改善反应时间具有明显促进作用。

- 脑疾病调控

迄今为止，国内外已开发出大量侵入性或非侵入性神经调控技术，主要包含电、磁、光、超声等。通过不同的刺激方式，神经调控可引发快速的、小范围的脑神经组织功能性改变，也可引起不间断的神经元功能改变以及神经环路连接改变[288]。作为诊断治疗神经系统疾病的有效手段，神经调控技术的临床应用价值已得到证实。

2.1 经颅磁刺激技术

经颅磁刺激是一种非侵入性的神经调控技术，通过时变磁场在大脑中产生感应电流，进而影响神经元的活动。常用的TMS常分为以下几类：

单脉冲经颅磁刺激（Single-pulse TMS, spTMS），是TMS技术的基础形式，通过向置于头皮的线圈通入瞬间强电流，产生短暂且高强度的磁场。该磁场能无阻碍地穿透头皮和颅骨，在大脑皮质中诱导出时变电场，可调节神经元活动，用于研究神经元功能。2022年Topka等人利用spTMS技术探究帕金森病（Parkinson’s disease, PD）患者在上肢运动冻结现象（ULF）期间运动皮层（M1）的兴奋性变化。研究结果表明帕金森病患M1兴奋性在冻结期间和冻结前的过渡期内降低，为进一步针对PD患者运动冻结症状的治疗提供了机制依据[289]。

成对脉冲刺激（Paired-pulse TMS, ppTMS），是指具有锁时关系的两个TMS脉冲，以特定时间间隔（Inter stimulus interval, ISI）成对出现，前一个脉冲对后一个脉冲产生影响，对大脑皮层产生不同效应的技术，第一个脉冲为条件刺激（Conditioning TMS, CS），第二个脉冲为实验刺激（Test TMS, TS）。ppTMS可用于评估健康人群及患者大脑皮层的抑制和易化机制，如帕金森病、肌张力障碍、阿尔茨海默病和精神分裂症等。2022的一项研究使用ppTMS来评估新兴技术经颅聚焦超声刺激对大脑可塑性研究的效果，通过测量SICI和ICF等指标，量化了对运动皮层内抑制和易化的影响，表明超声刺激可以调节皮层内抑制和兴奋的平衡，从而诱导人类运动皮层可塑性变化[290]。

重复经颅磁刺激（Repetitive TMS, rTMS），在单脉冲TMS基础上发展而来，通过在一段时间内重复发送磁脉冲，实现对大脑神经元活动的持续调节。其频率不同对皮层活动的影响存在差异，高频（High frequency rTMS, HF-rTMS）可增强皮层活动，低频（Low frequency rTMS, LF-rTMS）则抑制皮层活动，常用于治疗抑郁症等疾病。2021年的研究指出，可以使用rTMS对初级运动皮层M1施加10Hz的刺激，从而减缓周围神经性疼痛患者的头疼症状[291]；另外，rTMS还能成为对抗阿尔茨海默病的一种治疗策略，2022年Koch等人开展了为期24周的试验，以确定楔前叶刺激对轻度至中度阿尔茨海默病患者的安全性和有效性，结果表明24周的楔前叶rTMS能够减缓阿尔茨海默病患者的认知和功能下降，具有良好的应用前景[292]。

θ爆发式刺激（Theta burst stimulation, TBS），是rTMS的一种特殊形式，其刺激模式模拟了学习任务中大脑的自然活动，每20ms（50Hz）发送3个双相脉冲，间隔200ms（5Hz）。其中，间歇性TBS（iTBS）采用10s的刺激间隔，在2s的刺激时间内进行TBS模式刺激，能诱导长时程增强（LTP），增强突触连接强度，提升大脑皮质的兴奋性。连续TBS（cTBS）则是持续施加TBS脉冲20-40s，主要诱导长时程抑制（LTD），减弱突触连接强度，降低皮质兴奋性。TBS模式的优点在于能更高效地调节皮层兴奋性，且所需刺激时间更短、强度更低，在治疗神经系统及精神障碍疾病方面表现出良好的前景，如iTBS对抑郁症人群展现出良好的治疗潜力[293]、并在阿尔茨海默病的治疗中，通过增强神经可塑性，减少氧化应激，改善患者的认知功能[294]。

深部经颅磁刺激（dTMS）：dTMS旨在刺激大脑深部的神经组织。它通过特殊设计的线圈，产生能够深入大脑内部的磁场，突破了传统TMS主要刺激大脑浅层皮质的局限。这种技术利用了特殊的线圈结构和电流配置，使得磁场能够更有效地穿透到大脑深部区域，如前额叶深部、边缘系统等。这些深部脑区与情绪、认知等功能密切相关，dTMS对其进行刺激，为治疗一些复杂的神经精神疾病，如抑郁症、强迫症等提供了新的途径，有望更精准地调节相关神经回路，改善患者的症状[295-296]。

2.2 经颅电刺激技术

经颅电刺激是一种非侵入式的神经调控技术，通过在头皮施加微弱电流（通常不超过2mA）穿透颅骨至大脑皮层或深部神经元，进而调节细胞膜电位、神经振荡同步性或突触可塑性，实现对脑网络活动的定向干预[297]。因其操作简便、安全性高以及成本较低等优势，TES 已成为脑科学研究和临床治疗领域的重要工具之一。依据电流刺激方式分为经颅直流电刺激、经颅交流电刺激和经颅随机噪声刺激等[298]。

经颅直流电刺激（Transcranial Direct Current Stimulation, tDCS）通过施加恒定低强度直流电（1-2mA）实现皮层兴奋性的定向调控。阳极刺激通过持续去极化提高皮层神经元兴奋性，阴极刺激则引发超极化降低兴奋性。目前，tDCS在多种精神疾病的治疗中展现出良好的应用潜力。一项随机试验表明，阳极 tDCS刺激左侧前额叶皮层对抑郁症具有抗抑郁疗效，能够改善情绪并提升部分认知功能，且总体安全性和耐受性良好[299]。对于精神分裂症患者，阴极 tDCS 刺激左颞顶叶皮层联合阳极 tDCS 刺激左侧背外侧前额叶皮层（DLPFC），显著减少了幻觉[300]，并改善了阴性症状[301]。而一项meta分析研究显示，tDCS可以改善精神分裂症患者的注意力和工作记忆[302]。在神经退行性共济失调患者中，对双侧小脑皮质进行单次阳极 tDCS治疗可显著改善共济失调和步态[303]。此外，在一项tDCS与VR联合治疗创伤后应激障碍（Posttraumatic stress disorder, PTSD）的研究中，在1个月时tDCS真刺激组自我报告的PTSD症状严重程度显著降低，这表明tDCS联合VR可能是一种有前景的PTSD治疗策略[304]。

经颅交流电刺激（Transcranial Alternating Current Stimulation, tACS）通过在两个电极之间施加特定频率的微弱正弦电流来影响皮层节律，调节大脑兴奋性和自发电活动。多项研究表明，tACS可以显著提升认知功能。早期研究发现个性化α频率的tACS可以通过调节脑振荡活动来增强认知功能[305]。另一项研究强调了γ频率tACS在改善工作记忆方面的有效性，提示其在认知康复中的潜力[306]。2021年的一项随机对照试验表明，γ-tACS能够显著改善轻度认知障碍（MCI）和阿尔茨海默病（Alzheimer’ s disease, AD）患者的记忆表现，并可能通过增强胆碱能神经传导相关的皮层内连接性，促进认知功能的恢复[307]。2022 年，Zhou等人的研究进一步证实，40Hz的tACS能够显著改善阿尔茨海默病患者的认知功能，特别是在记忆、视觉和工作记忆方面[308]。此外，tACS在运动障碍和精神疾病中的应用也显示出积极效果。应用于运动皮层的 tACS可降低帕金森病患者的震颤幅度，实现静息震颤振幅平均降低近50%[309]。同时，针对强迫症患者眶额叶皮层的 tACS刺激能够有效缓解强迫症状[310]。

经颅随机噪声刺激（Transcranial Random Noise Stimulation, tRNS）使用随机强度的交流电，与tACS相比，tRNS具有更宽的频谱（0.1-640.0Hz）和随机噪声分布的特征，以覆盖大脑神经元振荡的频率范围。tRN能有效快速调节感知、改善学习，以及调节注意力、工作记忆等[311]。例如，在刺激阈值范围以内，使用tRNS将噪声集中添加到视觉皮层可以提高对阈下刺激的检测，同时将tRNS应用于其他皮质区域不会影响视觉检测准确性[312]。此外，在一项2020年的研究中发现，与tDCS和假刺激相比，tRNS下工作记忆得到了更明显且一致的增强[313]。

时间干涉电刺激（Temporal Interference Stimulation, TIS）是2017年Grossman在《Cell》杂志上提出了一种新型刺激方法，通过两束高频交流电信号（例如2000Hz和2010Hz）的干涉效应，在深部脑区形成低频包络信号（例如10Hz）。由于神经元的 “低通滤波性”，高频信号难以激活浅层皮层，而深部靶区可由低频包络信号有效刺激，从而实现避开浅层皮层刺激的同时刺激深部脑区[314]，解决了经颅磁刺激难以刺激深部、传统经颅电刺激难以聚焦深部的难题。研究表明，TIS 可以通过调节纹状体[315]和海马体[316]的记忆活动，增强运动学习和记忆表现。此外，TIS作为一种新型无创DBS方法，对于缓解轻度PD的运动症状是可行且安全的，尤其是运动迟缓和震颤[317]。而最近的研究显示，通过使用iTBS模式调节右海马-内嗅皮层的活动，可以改善空间导航表现[318]。时间干涉电刺激作为一种新兴的非侵入性脑刺激技术，在脑疾病的研究中展现出巨大的潜力。其高分辨率、高特异性和非侵入性的特点使其成为治疗神经退行性疾病、精神疾病和成瘾行为的有效手段。然而，目前的研究仍需进一步验证其在人类中的安全性和有效性，特别是在长期应用中的表现。

2.3 经颅超声刺激

无创经颅超声神经刺激（Transcranial Focused Ultrasound Stimulation, tFUS）作为一种新兴的脑疾病治疗手段，近年来在神经科学领域得到了广泛关注。用于神经调控应用的TUS通常指低强度聚焦超声刺激（Low-intensity transcranial Focused Ultrasound, LIFU），其主要特点是使用较低的声强（通常在mW/cm2到W/cm2范围内）发送超声波。相较于高强度聚焦超声（High-intensity Focused Ultrasound, HIFU），tFUS主要通过机械效应（如对离子通道产生影响、改变神经膜电位等）和空化效应而非热效应实现无创调控，不会引起明显的组织加热或不可逆损伤，因此可实现经颅无创刺激。相比之下，HIFU所使用的声场通常高达数百W/cm2，对组织造成的损伤通常不可逆，可实现热消融等用途[319]。

tFUS技术最显著的优势是具有较高的空间分辨率和较强的穿透力，可以在不损伤脑组织的情况下，通过声波精准地调控大脑深部区域的神经元活动。与经颅电刺激（tES）和经颅磁刺激（TMS）相比，tFUS可以以毫米级的空间分辨率靶向脑深部区域[320]。与DBS相比，FUS具有较小的侵入性，避免了手术风险，并可以针对不同靶区及功能进行连续和重复性的调控治疗。当前已有多项研究证实，tFUS可以靶向包括海马体、额叶、运动皮层、壳核和黑质在内的一系列浅层或深部的大脑区域[321-323]，通过重新定位或操纵换能器即可实现精准靶向。作为一种新的调控手段，tFUS的机制探究仍亟待揭示，临床应用也仍处于早期阶段，缺乏长期的安全性数据，但该种技术在多种疾病的调控中都展现出了较好的发展潜力和前景[324]。

在运动障碍领域，tFUS被应用于帕金森病PD和特发性震颤（ET）的治疗。2023年Samuel等人针对PD患者的研究表明，tFUS刺激双侧初级运动皮质可增加运动皮质兴奋性，是一种可行的PD非侵入性神经调节策略[325]。2024年Bancel等人针对特发性震颤患者的丘脑腹中间核（VIM）及齿状核-红核-丘脑束（DRT）进行了核磁影像导航的tFUS刺激，证明可以诱导震颤功率的持续大幅下降，一些患者的震颤功率降低约90%[326]。

在精神疾病方面，tFUS 通过调节大脑关键功能区域的神经活动，被探索用于治疗抑郁症（MDD）、焦虑症、强迫症（OCD）和精神分裂症等疾病，具有对情绪调节的潜在作用。例如，针对难治性抑郁症，部分研究表明患者在接受tFUS治疗后情绪改善可以增强自我报告的整体情绪并减少焦虑感、担忧等负面情绪[327]。Fan等人于2024年的研究表明tFUS通过作用于丘脑前核（ANT）可显著降低难治性抑郁症患者的抑郁评分（VAS-Depression和HAMD-6），同时减少默认模式网络（DMN）过度连接，这与抑郁症的神经病理机制相关[328]。

在神经退行性疾病方面，tFUS 主要被用于阿尔茨海默病（AD）和轻度认知障碍（MCI）的研究。tFUS可能通过作用于海马、前额叶皮层（PFC）等认知功能相关脑区，促进突触可塑性，提高认知功能。患者调查显示tFUS的治疗耐受性高，无主要副作用，其治疗后的神经心理评分显著改善，持续长达三个月，并且fMRI数据显示，症状的改善与记忆网络上调相关[329]。

在癫痫（Epilepsy）治疗方面，tFUS可作用于丘脑、颞叶或海马，抑制异常神经元放电，减少癫痫发作频率。临床试验研究表明，tFUS可安全地靶向癫痫患者的癫痫发作区（SOZ）[330]，降低癫痫患者脑电图EEG中的异常放电活动，部分患者在接受tFUS调控后癫痫发作减少，提示tFUS作为一种非侵入性干预手段的潜力[331]。

在慢性疼痛（Chronic Pain）管理方面，tFUS 被用于调控前扣带回、岛叶或丘脑的神经活动，以缓解疼痛感知。在一项随机交叉试验中，20名慢性疼痛患者针对前扣带回区域接受了40分钟的的真或假刺激。结果显示，FUS刺激在治疗后1天和7天内显著降低了疼痛强度，有效率达60%，而假刺激仅15%和20%有效。该技术安全、耐受性好，且效果持久，为慢性疼痛治疗提供了新的非侵入性选择[332]。

在物质使用障碍（Substance Use Disorder, SUD）的研究中tFUS也展现出良好的应用前景。2023年Mahoney等人报告了首例利用低强度聚焦超声靶向双侧伏隔核治疗SUD的人体研究。研究对象为一名43岁男性，患有严重阿片类、兴奋剂、尼古丁和大麻使用障碍。在接受治疗后，其对主要滥用物质的渴望显著降低（约50%），且在90天随访期内，对酒精和非法物质的渴望完全被抑制，尿毒理学检测也显示无药物使用[333]。该研究证实了LIFU的安全性和对SUD的潜在疗效，为未来治疗提供了新方向。

尽管无创超声调控技术在治疗脑疾病方面展现了巨大潜力，但仍面临一些挑战。首先，如何准确控制超声波的参数（如频率、强度、持续时间）以确保疗效同时避免副作用，是目前的研究热点。不同的脑区对超声的响应不同，如何实现个性化的治疗方案需要进一步探索。其次，超声的长期效应尚未完全明确，是否会对脑组织产生长期的影响仍需要更多的临床数据来验证[334]。

2.4 经颅光刺激技术

无创光生物调节技术，也被称为低强度光疗法（Low-level light therapy, LLLT）[335]，作为一种新兴的脑疾病治疗方法，近年来在神经科学领域获得了越来越多的关注。

光生物调节是利用红光（Red light，波长600-670nm）/近红外光（Near-infrared light，波长800-1100nm）波长范围的电磁波对大脑进行照射，该种电磁波可由低能量激光或发光二极管（LED）产生，可在体内及体外产生光生物调节效应[336]。其作用原理主要基于光与生物组织的相互作用。当特定波长的光（如红光或近红外光）照射生物组织时，光子被线粒体中的细胞色素c氧化酶（CCO）吸收，激活其电子传递链，从而增加ATP合成，提升细胞能量代谢。同时，光调节可促进钙离子释放，激活转录因子，调节基因表达，增强抗氧化防御，减轻炎症反应，并促进神经再生和突触形成。这些机制共同作用，改善细胞功能，促进组织修复，缓解疾病症状，尤其在神经退行性疾病和精神障碍中展现出显著的治疗潜力[337]。

光刺激的应用尤其在认知功能的改善与阿尔茨海默病AD的治疗中受到广泛关注。早在2016年，麻省理工学院蔡立慧团队就提出了“40Hz闪烁光”这种刺激模式的重要作用，证明该种刺激模式具有减少脑中的β淀粉样蛋白，改善小胶质细胞的重要功能，为AD的治疗提供了一种极具前景的调节手段[338]。

此后多个团队也利用光调控技术在人体中开展研究验证。例如，北师大牛海晶团队将1064nm的tPBM应用于右前额皮质（PFC），发现可以提高视觉工作记忆能力并增加枕顶叶对侧延迟活动（CDA）[339]。此后，针对健康老年人左侧背外侧前额叶进行一周光刺激（每天一次，每次12分钟），结果发现一周重复光刺激能够有效改善健康老年人工作记忆，并且这种改善效果能够维持至少3周[340]。2023年Lee等人综述tPBM在改善认知功能方面的数十项研究，表明tPBM可以改善健康成人的脑功能，并修复成人认知障碍的认知功能障碍，包括学习记忆、注意力、执行功能和语言功能[341]。

在帕金森病的治疗中，光生物调节同样表现出了有效性。tPBM治疗帕金森病的研究涵盖了动物实验和临床试验。动物模型中，tPBM显著减少了多巴胺能神经元的损伤，改善了运动功能障碍。临床试验也表明，tPBM能够缓解帕金森病患者的震颤、僵硬等症状，且未发现明显副作用。

除了神经退行性疾病，光生物调节在精神障碍的治疗中也展现了广泛应用前景。研究表明，通过特定波长（如810nm）的光刺激大脑，tPBM能够调节大脑神经活动、减轻炎症、促进神经可塑性，从而改善抑郁症患者的情绪状态和认知功能，缓解抑郁症状。类似地，光调节还可用于治疗创伤后应激障碍（PTSD）和强迫症等精神疾病[342]。

尽管无创光生物调节技术在治疗脑疾病方面展现了良好的前景，但仍面临一些挑战。首先，如何精确调控光的参数（如波长、光强、照射时间）以获得最佳治疗效果，是当前研究的难点。不同的脑区和病症对光的响应不同，如何根据患者的具体情况制定个性化的治疗方案需要进一步研究。其次，光照的长期效果尚未完全明确，需要更多的临床试验来验证其安全性和持续疗效。随着光源技术、传感器技术和脑影像技术的不断进步，未来无创光生物调节技术有望与其他治疗手段相结合，如神经调控技术、药物治疗等，形成综合治疗方案，进一步提高治疗效果。

- 运动辅助与康复

随着技术的发展，国际上面向运动功能辅助与康复的非侵入性脑机系统的研究已经取得了一系列重要成果。

2008年，美国得克萨斯大学奥斯汀分校研究团队开发了一种基于运动想象的非侵入BCI轮椅控制系统，通过解析左手运动想象、静息状态和特殊性单词想象产生的脑电信号，实现了轮椅的前进、左转和右转控制，首次证明了非侵入BCI在异步、连续控制外部设备中的可行性[343]。2014年，巴西科研团队在世界杯开幕式上展示了非侵入BCI与外骨骼机器人结合的创新应用，一名截瘫患者通过运动想象控制外骨骼完成开球动作，标志着非侵入BCI技术在复杂运动辅助中的突破性进展。2015年，意大利PERCRO实验室开发了基于运动想象的上肢外骨骼康复系统，通过实时解码患者运动意图并驱动机器人辅助抓握动作，验证了闭环康复训练的有效性[344]。2018年，瑞士洛桑联邦理工学院团队提出脑机接口联合功能性电刺激（BCI-FES）的治疗方案，通过随机对照试验发现，实验组患者运动功能恢复与神经功能连接增强显著相关，揭示了BCI技术通过激活中枢神经传出与感觉反馈通路促进功能重塑的机制[345]。2019年德国研究团队通过随机对照试验发现，基于运动想象的脑机接口（MI-BCI）训练可显著改善脑卒中患者上肢运动功能（FMA-UE评分提高），并诱导受损脑区皮质功能重组[346]。2022年，瑞士苏黎世大学团队采用经颅磁刺激联合脑电技术（TMS-EEG）评估卒中患者运动皮层兴奋性，发现间歇性θ波爆发刺激（iTBS）可降低双侧运动皮层N100振幅，为神经调控疗效提供了电生理证据[347]。2023年，俄罗斯科学院高级神经活动和神经生理学研究所团队将功能性近红外光谱BCI应用于卒中后手功能康复，通过运动想象训练实现了短期运动功能改善，验证了多模态脑信号采集技术的临床价值[348]。2024年，丹麦奥胡斯大学团队研发了低成本开源BCI系统与3D打印腕部外骨骼，通过家庭环境适应性验证，证实该系统可有效诱导神经可塑性并促进运动功能恢复[349]。2024年美国斯坦福大学团队开发了通用智能BCI系统NOIR，利用非侵入式脑电技术实现人类对烹饪、清洁等20项日常任务的机器人控制，展示了非侵入BCI在复杂环境中的应用潜力[350]。

近年来，我国在非侵入BCI系统在运动功能辅助与康复领域的研究也取得多项重大进展，部分成果已进入临床验证与产业化阶段。2015年，天津大学团队研发的“神工”系列人工神经康复机器人系统，融合运动想象疗法与多级神经肌肉电刺激技术，在体外仿生构建“中枢-外周”神经通路，临床试验显示卒中患者康复有效率超过70%。2016年，复旦大学团队对脑卒中患者进行闭环MI-BCI康复训练，患者在平衡量表评分标准表（Fugl-Meyer Assessment, FMA）、运动功能状态量表（Motor status scale, MSS）评分上较治疗前均有所提高[351]。2020年，清华大学团队对中重度上肢功能障碍的亚急性脑卒中患者进行4周的综合康复训练，包括MI-BCI、常规运动疗法和作业疗法，训练后患者FMA、手臂动作调查测试表（Action Research Arm Test , ARAT）和Wolf运动功能测试（Wolf Motor Function Test, WMFT）评分均明显升高，多个脑区间功能连接增强[352]。2022年，浙江大学团队提出了一种通过基于稳态视觉诱发电位的脑机接口控制的软体机器人手套，用于中风后手功能康复的研究提高了康复效果，增强了患者的主动参与感和康复的积极性[353]。同年，强脑科技推出BrainRobotics智能仿生手，通过深度学习算法解析表面肌电信号，实现上肢截肢患者对复杂手部动作的思维控制[354]。2024年，北京航空航天大学团队开发虚拟现实（VR）增强型运动想象范式，通过第三人称视角的虚拟化身体感诱导，使脑卒中患者患侧感觉运动区事件相关去同步化（ERD）水平显著提升，增强了运动想象任务的大脑响应特征[355]。同年，天津大学团队提出两项创新技术：一是“第六指”运动想象增强范式，利用外源性肢体重塑大脑映射表征，使感觉运动区ERD水平提升188%，前额叶区域提升117%[356-357]；二是步态序列编码的体感电刺激引导系统，通过模拟真实行走的肌群时序激活模式，将下肢运动想象分类准确率提升至80%以上[358]。

# 参考文献

- Tang X H, Shen H, Zhao S, et al. Flexible brain–computer interfaces[J]. Nature Electronics, 2023, 6:109-118.

- Moon H, Kwon J, Eun J, et al. Electrocorticogram (ECoG): Engineering Approaches and Clinical Challenges for Translational Medicine[J]. Advanced Materials Technologies, 2024, 9(12).

- Zhang A, Mandeville E T, et al. Ultraflexible endovascular probes for brain recording through micrometer-scale vasculature[J]. Science, 2023, 381: 306-312.

- Wang X, Wu S, Yang H, et al. Intravascular delivery of an ultraflexible neural electrode array for recordings of cortical spiking activity[J]. Nature Communications, 2024, 15:9442.

- Wang J, Wang T, Liu H, et al. Flexible electrodes for brain–computer interface system[J]. Advanced materials, 2023:2211012.

- Luan L, Wei X, Zhao Z, et al. Ultraflexible nanoelectronic probes form reliable, glial scar–free neural integration[J]. Science Advances, 2017, 3(2):e1601966.

- Zhao Z, Zhu H, Li X, et al. Ultraflexible electrode arrays for months-long high-density electrophysiological mapping of thousands of neurons in rodents[J]. Nature Biomedical Engineering, 2022.

- Guan S, Wang J, Gu X, et al. Elastocapillary self-assembled neurotassels for stable neural activity recordings[J]. Science advances, 2019, 5(3):eaav2842.

- Liu Y, Jia H, Sun H, et al. A high-density 1,024-channel probe for brain-wide recordings in non-human primates[J]. Nature Neuroscience, 2024, 27:1620-1631.

- Wang S, Jiang Q, Liu H, et al. Mechanically adaptive and deployable intracortical probes enable long-term neural electrophysiological recordings[J]. Proceedings of the National Academy of Sciences, 2024, 121(40):e2403380121.

- Musk E. An Integrated Brain-Machine Interface Platform With Thousands of Channels[J]. Journal of Medical Internet Research, 2019, 21(10):e16194.

- Yoon D Y, Pinto S, Chung S W, et al. A 1024-channel simultaneous recording neural SoC with stimulation and real-time spike detection[C]. Symposium on VLSI Circuits IEEE, 2021.

- Ward S, Riley C, Carey E M, et al. Electro-optical mechanically flexible coaxial microprobes for minimally invasive interfacing with intrinsic neural circuits[J]. Nature Communications, 2022, 13.

- Park S, Yuk H, Zhao R, et al. Adaptive and multifunctional hydrogel hybrid probes for long-term sensing and modulation of neural activity[J]. Nature communications, 2021, 12(1):3435.

- Yang X, Zhou T, Zwang T J, et al. Bioinspired neuron-like electronics[J]. Nature Materials, 2019:510-517.

- Li Q, Nan K, Floch P L, et al. Cyborg organoids: implantation of nanoelectronics via organogenesis for tissue-wide electrophysiology[J]. Nano Letters, 2019, 19(8):5781-5789.

- Kim S, Bhandari R, Klein M, et al. Integrated wireless neural interface based on the Utah electrode array[J]. Biomedical Microdevices, 2009, 11(2):453–466.

- An H, et al. A power-efficient brain-machine interface system with a sub-mW feature extraction and decoding ASIC demonstrated in nonhuman primates[J]. IEEE Transactions on Biomedical Circuits and Systems, 2022, 16(3):395–408.

- Lee J, Leung V, Lee A H, et al. Neural recording and stimulation using wireless networks of microimplants[J]. Nature Electronics, 2021, 4(8):604–614.

- Obaid A, Hanna M E, Wu Y W, et al. Massively parallel microwire arrays integrated with CMOS chips for neural recording[J]. Science Advances, 2020, 6(12):eaay2789.

- Smith W A, Uehlin J P, Perlmutter S I, et al. A scalable, highly-multiplexed delta-encoded digital feedback ECoG recording amplifier with common and differential-mode artifact suppression[C]. Symposium on VLSI Circuits IEEE, 2017.

- Aymerich J, Sawigun C, Cisneros-Fernandez J, et al. A 16-Ch CMI-Tolerant Neural AFE with Inherent CM Detection and Shared CM Suppression Achieving 0.006 mm²/Ch and 3.1 μW/Ch[C]. IEEE Symposium on VLSI Technology and Circuits, 2024.

- Intan Technologies. RHD2164 Datasheet. Retrieved from https://intantech.com/files/Intan_RHD2164_datasheet.pdf, 2024.

- Intan Technologies. RHS2116 Datasheet. Retrieved from https://intantech.com/files/Intan_RHS2116_datasheet.pdf, 2024.

- Neuropixels. Neuropixels 2.0 Datasheet. Retrieved from https://neuropixels.org/, 2024.

- Neuralink. Neuralink Whitepaper. Retrieved from https://www.documentcloud.org/documents/6204648-Neuralink-White-Paper/, 2024

- Blackrock Neurotech. Blackrock Neurotech Datasheet. Retrieved from https://blackrockneurotech.com/products/, 2024

- NeuroNexus. SiNAPS Pixel Probes Datasheet. Retrieved from https://www.neuronexus.com/products/electrode-arrays/sinaps/, 2024.

- 神芯科技. 128通道采集芯片产品手册. Retrieved from https://sensingx.com.cn/support/, 2025.

- 矩智科技. MTXNEP256 Datasheet. Retrieved from https://mtrix.cn/newsinfo/8195106.html, 2024.

- 微灵医疗. 硬膜外神经信号采集芯片. Retrieved from https://we-linking.com/newsinfo/7919408.html, 2024.

- Joshua E. Woods，et al. Miniature battery-free epidural cortical stimulators[J]. Science advances, 2024, 10(15): eadn0858.

- Xinmiao Zhang, Jiawei Li, Zhuoran Li, Bo Hong, et al. Leading and following: Noise differently affects semantic and acoustic processing during naturalistic speech comprehension[J]. NeuroImage, 2023, 282:120404.

- 澳大利亚compumedics neuroscan公司. Retrieved from https://compumedicsneuroscan.com.

- 奥地利g.Tec公司. Retrieved from https://www.gtec.at.

- 美国ceribell公司. Retrieved from https://ceribell.com/.

- 浙江柔灵科技有限公司. Retrieved from https://www.flexolinkai.com/index.html

- Pei W, Wu X, Zhang X, et al. A pre-gelled eeg electrode and its application in ssvep-based bci[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2022, 30:843-850.

- Ahmed S, Momin M, Ren J, et al. Stick-and-play bioadhesive hairlike electrodes for chronic EEG recording on human[J]. npj Biomedical Innovations, 2025, 2(1):9.

- Tian L, Zimmerman B, Akhtar A, et al. Large-area mri-compatible epidermal electronic interfaces for prosthetic control and cognitive monitoring[J]. Nature biomedical engineering, 2019, 3(3):194-205.

- Tian Q, Zhao H, Wang X, et al. Hairy-skin-adaptive viscoelastic dry electrodes for long-term electrophysiological monitoring[J]. Advanced Materials, 2023, 35(30):2211236.

- Liu Z, Wang Y, Pei W, et al. A high-resolution dry electrode array for ssvep-based brain-computer interfaces[C]. Proceedings of the 2019 9th International IEEE/EMBS Conference on Neural Engineering. San Francisco, USA: IEEE, 2019:811-814.

- Erickson B, Rich R, Shankar S, et al. Evaluating and benchmarking the EEG signal quality of high-density, dry mxene-based electrode arrays against gelled ag/agcl electrodes[J]. Journal of Neural Engineering, 2024, 21(1):016005.

- Yuan H, Li Y, Yang J, et al. State of the Art of Non-Invasive Electrode Materials for Brain–Computer Interface[J]. Micromachines, 2021, 12(12):1521.

- Mahmood M, Kwon S, Kim H, et al. Wireless soft scalp electronics and virtual reality system for motor imagery-based brain-machine interfaces[J]. Advanced Science, 2021, 8(19):2101129.

- Kim H, Kim J H, Lee Y J, et al. Motion artifact–controlled micro–brain sensors between hair follicles for persistent augmented reality brain–computer interfaces[J]. Proceedings of the National Academy of Sciences, 2025, 122(15):e2419304122.

- Li J, Ma Y, Huang D, et al. High-performance flexible microneedle array as a low-impedance surface biopotential dry electrode for wearable electrophysiological recording and polysomnography[J]. Nano-Micro Letters, 2022, 14(1):132.

- Zhang J, Zhao Y, Li J, et al. Stretchable microneedle array electrode for movement-resistant electroencephalogram monitoring[C]. Proceedings of the 2025 26th International Conference on Solid-State Sensors, Actuators and Microsystems. Orlando, USA: IEEE, 2019.

- Liu Z, Mei J, Tang J, et al. A memristor-based adaptive neuromorphic decoder for brain–computer interfaces[J]. Nature Electronics‌, 2025, 8:362–372.

- Uran A, et al. A 16-Channel Neural Recording System-on-Chip With CHT Feature Extraction Processor in 65-nm CMOS[J]. IEEE Journal of Solid-State Circuits, 2022, 57(9): 2752-2763.

- Kaifosh P, Reardon T R, Skipper M. A generic non-invasive neuromotor interface for human-computer interaction[J]. Nature, 2025.

- Mei J, Luo R, Xu L, et al. MetaBCI: An open-source platform for brain–computer interfaces[J]. Computers in Biology and Medicine, 2024, 168:107806.

- Digiovanna J, Mahmoudi B, Fortes J, et al. Coadaptive brain-machine interface via reinforcement learning[J]. IEEE transactions on biomedical engineering, 2009, 56(1):54-64.

- Wang Y, Wang F, Xu K, et al. Neural control of a tracking task via attention-gated reinforcement learning for brain-machine interfaces[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2015, 23(3):458-467.

- Wang F, Wang Y, Xu K, et al. Quantized Attention-Gated Kernel Reinforcement Learning for Brain-Machine Interface Decoding[J]. IEEE transactions on neural networks and learning systems, 2015, 28(4):873-886.

- Zhang X, Libedinsky C, So R, et al. Clustering Neural Patterns in Kernel Reinforcement Learning Assists Fast Brain Control in Brain-Machine Interfaces[J]. IEEE transactions on neural systems and rehabilitation engineering, 2019, 27(9):1684-1694.

- X Zhang, S Chen, Y Wang. Kernel Reinforcement Learning-Assisted Adaptive Decoder Facilitates Stable and Continuous Brain Control Tasks[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2023, 31:4125-4134.

- J Bae, L G. et al. Kernel temporal differences for neural decoding[J]. Computational intelligence and neuroscience, 2015, 1:481375.

- Bae J, Chhatbar P, Francis J T, et al. Reinforcement learning via kernel temporal difference[J]. International Conference of the IEEE Engineering in Medicine & Biology Society, 2011:5662-5665.

- Bae J, Giraldo L G S, Francis J T, et al. Correntropy kernel temporal differences for reinforcement learning brain machine interfaces[C]. International Joint Conference on Neural Networks (IJCNN), 2014:2713-2717.

- X Shen, X Zhang, Y Huang, et al. Intermediate Sensory Feedback Assisted Multi-Step Neural Decoding for Reinforcement Learning based Brain-Machine Interfaces[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2022:1-1.

- Mahmoudi B, Pohlmeyer E A, Prins N W, et al. Towards autonomous neuroprosthetic control using Hebbian reinforcement learning[J]. Journal of Neural Engineering, 2013, 10(6):066005.

- Pohlmeyer E A, Mahmoudi B, Geng S, et al. Brain-machine interface control of a robot arm using actor-critic rainforcement learning[C]. Annual International Conference of the IEEE Engineering in Medicine and Biology Society, 2012:4108-4111.

- Mahmoudi B, Sanchez J C. A Symbiotic Brain-Machine Interface through Value-Based Decision Making[J]. Plos One, 2011, 6(3):1-14.

- J Tan, X Zhang, et al. Hidden Brain State-based Internal Evaluation Using Kernel Inverse Reinforcement Learning in Brain-machine Interfaces[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2024.

- Zhang X, Shen X, Wang Y. Modeling Task Engagement to Regulate Reinforcement Learning-based Decoding for Online Brain Control[J]. IEEE Transactions on Cognitive and Developmental Systems, 2024.

- Zhang X, Wang Y. Covariant Cluster Transfer for Kernel Reinforcement Learning in Brain-Machine Interface*[C]. Annual International Conference of the IEEE Engineering in Medicine Biology Society (EMBC), 2020:3086-3089.

- Zhang X, Wang Y. A Weight Transfer Mechanism for Kernel Reinforcement Learning Decoding in Brain-Machine Interfaces[C]. International Conference of the IEEE Engineering in Medicine & Biology Society (EMBC), 2019.

- L Paninski, Maximum likelihood estimation of cascade point-process neural encoding models[J]. Network: Computation in Neural Systems, 2004, 15(4):243–262.

- Wang Y, Principe J C. Instantaneous estimation of motor cortical neural encoding for online brain–machine interfaces[J].  Journal Of Neural Engineering, 2010, 7(5):056010.

- Shuhang Chen, Xiang Zhang, et al. Tracking Fast Neural Adaptation by Globally Adaptive Point Process Estimation for Brain-Machine Interface[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2021, 29:1690–1700.

- S Kim, D Putrino, et al. A Granger Causality Measure for Point Process Models of Ensemble Neural Spiking Activity[J]. PLoS Computational Biology, 2011, 7(3):e1001110.

- Casile A, Faghih R T, Brown E N. Robust point-process Granger causality analysis in presence of exogenous temporal modulations and trial-by-trial variability in spike trains[J]. PLOS Computational Biology, 2021, 17(1):e1007675.

- Sani O G, Abbaspourazad H, Wong Y T, et al. Modeling behaviorally relevant neural dynamics enabled by preferential subspace identification[J]. Nature Neuroscience, 2021, 24(1):140–149.

- Li M, Chen S, Zhang X, et al. Neural Correlation Integrated Adaptive Point Process Filtering on Population Spike Trains[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2025, 33:1014–1025.

- Gallego J A, Perich M G, Chowdhury R H, et al. Long-term stability of cortical population dynamics underlying consistent behavior[J]. Nature Neuroscience, 2020, 23(2):260–270.

- Abbaspourazad H, Erturk E, Pesaran B, et al. Dynamical flexible inference of nonlinear latent factors and structures in neural population activity[J]. Nature biomedical engineering, 2025, 2:9.

- Depasquale B, Sussillo D, Abbott L F, et al. The centrality of population-level factors to network computation is demonstrated by a versatile approach for training spiking networks[J]. Neuron, 2023, 111(5):631-649.e10.

- Wang Y, She X, Liao Y, et al. Tracking Neural Modulation Depth by Dual Sequential Monte Carlo Estimation on Point Processes for Brain–Machine Interfaces[J]. IEEE Transactions on Biomedical Engineering, 2016, 63(8):1728–1741.

- S Chen, X Zhang, et al. Online Estimating Pairwise Neuronal Functional Connectivity in Brain–Machine Interface[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering. 2024, 32:271–281.

- Capogrosso M. A brain-spine interface alleviating gait deficits after spinal cord injury in primates[J]. Nature, 2016, 539(7628):284–288.

- F B Wagner, et al. Targeted neurotechnology restores walking in humans with spinal cord injury[J]. Nature, 2018, 563(7729):65–71.

- Hankov N, Caban M, Demesmaeker R, et al. Augmenting rehabilitation robotics with spinal cord neuromodulation: a proof of concept[J]. Science Robotics, 2025, 10(100).

- Deadwyler S A, Hampson R E, Song D, et al. A cognitive prosthesis for memory facilitation by closed-loop functional ensemble stimulation of hippocampal neurons in primate brain[J]. Experimental Neurology, 2016, 287:452-460.

- Hampson R E, Song D, Robinson B S, et al. Developing a hippocampal neural prosthetic to facilitate human memory encoding and recall[J]. Journal of Neural Engineering, 2018, 15(3):036014.

- Roeder B M, et al. Developing a hippocampal neural prosthetic to facilitate human memory encoding and recall of stimulus features and categories[J]. Frontiers in Computational Neuroscience, 2024, 18:1263311.

- Song D, Robinson B, Hampson R, et al. Sparse Large-Scale Nonlinear Dynamical Modeling of Human Hippocampus for Memory Prostheses[J]. IEEE Transactions on Neural Systems & Rehabilitation Engineering A Publication of the IEEE Engineering in Medicine & Biology Society, 2016, 26(2):272–280.

- Truccolo W. A Point Process Framework for Relating Neural Spiking Activity to Spiking History, Neural Ensemble, and Extrinsic Covariate Effects[J]. Journal of Neurophysiology, 2005, 93(2):1074-1089.

- Cunle, Qian, Xuyun, et al. Nonlinear Modeling of Neural Interaction for Spike Prediction Using the Staged Point-Process Model.[J]. Neural Computation, 2018, 30(12):3189–3226.

- Wu S, et al. Generative neural spike prediction from upstream neural activity via behavioral reinforcement. bioRxiv. doi: 10.1101/2023.07.25.550495, 2023.

- Wu S, Zhang X, Wang Y. Neural Manifold Constraint for Spike Prediction Models Under Behavioral Reinforcement[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2024, 32:2772–2781.

- Schirrmeister R T, Springenberg J T, Fiederer L D J, et al. Deep learning with convolutional neural networks for EEG decoding and visualization[J]. Human Brain Mapping, 2017, 38(11):5391-5420.

- Song Y H, et al. EEGConformer: Convolutional transformer for EEG decoding and visualization[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2022, 31:710–719.

- Jiang W B, Zhao L M, Lu B L. Large brain model for learning generic representations with tremendous EEG data in BCI[C]. The International Conference on Learning Representations. Vienna, Austria, 2024.

- Wang G, Liu W, He Y, et al. EEGPT: Pretrained Transformer for Universal and Reliable Representation of EEG Signals[C].  Proceedings of the Neural Information Processing Systems Conference. New Orleans, USA, 2025.

- Banville H, Chehab O, Hyvärinen A, et al. Uncovering the structure of clinical EEG signals with self-supervised learning[J]. Journal of Neural Engineering, 2021, 18(4):046020.

- Kostas D, Aroca-Ouellette S, Rudzicz F. BENDR: Using transformers and a contrastive self-supervised learning task to learn from massive amounts of EEG data[J]. Frontiers in Human Neuroscience, 2021, 15:653659.

- Chien H Y S, Goh H, Sandino C M, et al. MAEEG: Masked auto-encoder for EEG representation learning. arXiv:2211.02625, 2022.

- Wang C, Subramaniam V, Yaari A U, et al. BrainBERT: Self-supervised representation learning for intracranial recordings[C]. The International Conference on Learning Representations, 2022.

- Zhang D, Yuan Z, Yang Y, et al. Brant: Foundation model for intracranial neural signal[C]. Advances in Neural Information Processing Systems. New Orleans, USA, 2023.

- Yuan Z, Zhang D, Chen J, et al. Brant-2: Foundation model for brain signals. arXiv:2402.10251, 2024.

- Zhang D, Yuan Z, Chen J, et al. Brant-X: A unified physiological signal alignment framework[C]. Proceedings of the 30th ACM SIGKDD Conference on Knowledge Discovery and Data Mining. Singapore, 2024:4155-4166.

- Yang Y, Li J, Zhang D, et al. BIOT: A generic biosignal learning model for joint pre-training and knowledge transfer across datasets. arXiv:2309.12345, 2023.

- Foumani N M, Mackellar G, Ghane S, et al. EEG2Rep: Enhancing self-supervised EEG representation through informative masked inputs. arXiv:2402.17772, 2024.

- Wan Z, Yang R, Huang M, et al. A review on transfer learning in EEG signal analysis[J]. Neurocomputing, 2021, 421:1-14.

- H He, D Wu. Transfer learning for brain-computer interfaces: A Euclidean space data alignment approach[J]. IEEE Trans. on Biomedical Engineering, 2020, 67(2):399–410.

- Li S, Wang Z, Luo H, et al. T-TIME: Test-time information maximization ensemble for plug-and-play BCIs[J]. IEEE Trans. on Biomedical Engineering, 2024, 71(2):423–432.

- K Xia, et al. Privacy-Preserving Brain-Computer Interfaces: A Systematic Review[J]. IEEE Transactions on Computational Social Systems, 2023, 10(5):2312-2324.

- D'Anna E, Valle G, Mazzoni A, et al. A closed-loop hand prosthesis with simultaneous intraneural tactile and position feedback[J]. Science Robotics, 2018, 4(27):eaau8892.

- Ganzer P D, Colachis S C, Schwemmer M A, et al. Restoring the Sense of Touch Using a Sensorimotor Demultiplexing Neural Interface[J]. Cell, 2020, 181(4):763-773.e12.

- 马林, 巢宝华, 曹雷等. 2007—2017年中国脑卒中流行趋势及特征分析[J]. 中华脑血管病杂志(电子版), 2020, 14(05):253.

- Maas A I, Stocchetti N, Bullock R. Moderate and severe traumatic brain injury in adults[J]. The Lancet Neurology, 2008, 7(8):728-741.

- Feigin V L, Nichols E, Alam T, et al. Global, regional, and national burden of neurological disorders, 1990–2016: a systematic analysis for the Global Burden of Disease Study 2016[J]. The Lancet Neurology, 2019, 18(5):459-480.

- Penfield W, Boldrey E. Somatic motor and sensory representation in the cerebral cortex of man as studied by electrical stimulation[J]. Brain, 1937, 60(4):389-443.

- Chang E F. Towards large-scale, human-based, mesoscopic neurotechnologies[J]. Neuron, 2015, 86(1):68-78.

- Nai D, Simon J Z. Cortical entrainment to continuous speech: functional roles and interpretations[J]. Frontiers in human neuroscience, 2014, 8:311.

- Ding N, Melloni L, Zhang H, et al. Cortical tracking of hierarchical linguistic structures in connected speech[J]. Nature neuroscience, 2016, 19(1):158-164.

- Tourville J A, Guenther F H. The DIVA model: A neural theory of speech acquisition and production[J]. Language and cognitive processes, 2011, 26(7):952-981.

- Crone N E, Hao L, Hart J, et al. Electrocorticographic gamma activity during word production in spoken and sign language[J]. Neurology, 2001, 57(11): 2045-2053.

- Bouchard K E, Mesgarani N, Johnson K, et al. Functional organization of human sensorimotor cortex for speech articulation[J]. Nature, 2013, 495(7441): 327-332.

- Bouchard K E, Mesgarani N, Johnson K, et al. Encoding of articulatory kinematic trajectories in human speech sensorimotor cortex[J]. Neuron, 2018, 98(5): 1042-1054. e4.

- Dichter B K, Breshears J D, Leonard M K, et al. The control of vocal pitch in human laryngeal motor cortex[J]. Cell, 2018, 174(1):21-31.e9.

- Li Y, Tang C, Lu J, et al. Human cortical encoding of pitch in tonal and non-tonal languages[J]. Nature Communications, 2021, 12(1):1161.

- Lu J, Li Y, Zhao Z, et al. Neural control of lexical tone production in human laryngeal motor cortex[J]. Nature Communications, 2023, 14(1):6917.

- Qi Y, Zhu X, Xiong X, et al. Human motor cortex encodes complex handwriting through a sequence of stable neural states[J]. Nature Human Behaviour, 2025, 9(6):1260-1271.

- Rabbani Q, Milsap G, Crone N E. The potential for a speech brain–computer interface using chronic electrocorticography[J]. Neurotherapeutics, 2019, 16(1):144-165.

- Brandman D M, Cash S S, Hochberg L R. human intracortical recording and neural decoding for brain–computer interfaces[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2017, 25(10):1687-1696.

- Wilson G H, Stavisky S D, Willett F R, et al. Decoding spoken English from intracortical electrode arrays in dorsal precentral gyrus[J]. Journal of neural engineering, 2020, 17(6):066007.

- Wilson G H, Stavisky S D, Willett F R, et al. Direct classification of all American English phonemes using signals from functional speech motor cortex[J]. Journal of neural engineering, 2014, 11(3):035015.

- Brumberg J S, Wright E J, Andreasen D S, et al. Classification of intended phoneme production from chronic intracortical microelectrode recordings in speech motor cortex[J]. Frontiers in neuroscience, 2011, 5:7880.

- Ramsey N F, Salari E, Aarnoutse E J, et al. Decoding spoken phonemes from sensorimotor cortex with high-density ECoG grids[J]. Neuroimage, 2018, 180:301-311.

- Herff C, Heger D, Pesters A D, et al. Brain-to-text: decoding spoken phrases from phone representations in the brain[J]. Frontiers in neuroscience, 2015, 9:217.

- Kellis S, Miller K, Thomson K, et al. Decoding spoken words using local field potentials recorded from the cortical surface[J]. Journal of neural engineering, 2010, 7(5):056007.

- Livezey J A, Bouchard K E, Chang E F. Deep learning as a tool for neural data analysis: speech classification and cross-frequency coupling in human sensorimotor cortex[J]. PLoS computational biology, 2019, 15(9):e1007091.

- Wairagkar M, Hochberg L R, Brandman D M, et al. Synthesizing speech by decoding intracortical neural activity from dorsal motor cortex[C]. 2023 11th International IEEE/EMBS Conference on Neural Engineering (NER). IEEE, 2023: 1-4.

- Li J, Guo C, Fu L, et al. Neural2Speech: A transfer learning framework for neural-driven speech reconstruction[C]. ICASSP 2024-2024 IEEE International Conference on Acoustics, Speech and Signal Processing (ICASSP). IEEE, 2024:2200-2204.

- Liu Y, Zhao Z, Xu M, et al. Decoding and synthesizing tonal language speech from brain activity[J]. Science Advances, 2023, 9(23):eadh0478.

- Anumanchipalli G K, Chartier J, Chang E F. Speech synthesis from neural decoding of spoken sentences[J]. Nature, 2019, 568(7753):493-498.

- Zhang Daohan, Wang Zhenjie, Qian Youkun, et al. A brain-to-text framework for decoding natural tonal sentences[J]. Cell Reports, 2024, 43(11).

- Feng Chen, Cao Lu, Wu Di, et al. Acoustic inspired brain-to-sentence decoder for logosyllabic language[J]. Cyborg and Bionic Systems, 2025(6).

- Wu Jinsong, Zhou Zhitao, Qi Jianxuan, et al. Real-time decoding of full-spectrum Chinese using brain-computer interface[J]. Science Advances, 2025, 11(45):eadz9968.

- Roussel P, Godais G L, Bocquelet F, et al. Observation and assessment of acoustic contamination of electrophysiological brain signals during speech production and sound perception[J]. Journal of Neural Engineering, 2020, 17(5):056028.

- Moses D A, Leonard M K, Makin J G, et al. Real-time decoding of question-and-answer speech dialogue using human cortical activity[J]. Nature Communications, 2019, 10(1):3096.

- Moses D A, Metzger S L, Liu J R, et al. Neuroprosthesis for decoding speech in a paralyzed person with anarthria[J]. New England Journal of Medicine, 2021, 385(3):217-227.

- Metzger S L, Liu J R, Moses D A, et al. Generalizable spelling using a speech neuroprosthesis in an individual with severe limb and vocal paralysis[J]. Nature Communications, 2022, 13(1):6510.

- Silva A B, Liu J R, Metzger S L, et al. A bilingual speech neuroprosthesis driven by cortical articulatory representations shared between languages[J]. Nature Biomedical Engineering, 2024, 8(8):977-991.

- Willett F R, Kunz E M, Fan C, et al. A high-performance speech neuroprosthesis[J]. Nature, 2023, 620(7976):1031-1036.

- Metger S L, Littlejohn K T, Silva A B, et al. A high-performance neuroprosthesis for speech decoding and avatar control[J]. Nature, 2023, 620(7976):1037-1046.

- Kunz E M, Krasa B A, Kamdar F, et al. Inner speech in motor cortex and implications for speech neuroprostheses[J]. Cell, 2025, 188(17):4658-4673.e17.

- Card N S, Wairagkar M, Iacobacci C, et al. An accurate and rapidly calibrating speech neuroprosthesis[J]. New England Journal of Medicine, 2024, 391(7):609-618.

- Littlejohn K T, Cho C J, Liu J R, et al. A streaming brain-to-voice neuroprosthesis to restore naturalistic communication[J]. Nature Neuroscience, 2025, 28(4):902-912.

- Zrenner E. Will retinal implants restore vision?[J]. Science, 2002, 295(5557):1022-1025.

- Santos A, et al. Preservation of the inner retina in retinitis pigmentosa: A morphometric analysis[J]. Archives of Ophthalmology, 1997, 115(4):511-515.

- Ozmert E, Arslan U. Retinal Prostheses and Artificial Vision[J]. Retinal Prostheses and Artificial Vision, 2019, 49:213-219.

- Chuang A T, Margo C E, Greenberg P B. Retinal implants: a systematic review[J]. British Journal of Ophthalmology, 2014, 98(7):852-856.

- Ho A C, et al. Long-Term Results from an Epiretinal Prosthesis to Restore Sight to the Blind[J]. Ophthalmology, 2015, 122(8):1547-155.

- da Cruz L, et al. Five-Year Safety and Performance Results from the Argus II Retinal Prosthesis System Clinical Trial[J]. Ophthalmology, 2016, 123(10):2248-2254.

- Christie B, et al. Sequential epiretinal stimulation improves discrimination in simple shape discrimination tasks only[J]. Journal of Neural Engineering, 2022, 19.

- Beyeler M, et al. A model of ganglion axon pathways accounts for percepts elicited by retinal implants[J]. Cold Spring Harbor Laboratory, 2018, 1.

- Muqit M M K, et al. Six-Month Safety and Efficacy of the Intelligent Retinal Implant System II Device in Retinitis Pigmentosa[J]. Ophthalmology, 2018.

- Keseru M. et al. Acute electrical stimulation of the human retina with an epiretinal electrode array[J]. Acta Ophthalmologica, 2012, 90(1):e1-8.

- Stingl K, et al. Functional outcome in subretinal electronic implants depends on foveal eccentricity[J]. Investigative Ophthalmology & Visual Science, 2013, 54(12):7658-7665.

- Stingl K. et al. Artificial vision with wirelessly powered subretinal electronic implant alpha-IMS[J]. Proceedings of the Royal Society B: Biological Sciences, 2013.

- Stingl K. et al. Subretinal Visual Implant Alpha IMS--Clinical trial interim report[J]. Vision Research, 2015.

- Stingl, K. et al. Interim Results of a Multicenter Trial with the New Electronic Subretinal Implant Alpha AMS in 15 Patients Blind from Inherited Retinal Degenerations[J]. Frontiers in Neuroscience, 2017, 11:445.

- Edwards T L, et al. Assessment of the Electronic Retinal Implant Alpha AMS in Restoring Vision to Blind Patients with End-Stage Retinitis Pigmentosa[J]. Ophthalmology, 2018, 125(3):432-443.

- Yanovitch L, Raz-Prag D, Hanein Y. A new high-resolution three-dimensional retinal implant: System design and preliminary human results. bioRxiv:doi:10.1101/2022.09.14.50790, 2022.

- Pixium Vision announces FDA approval of clinical trial expansion of Prima System US feasibility study to Stanford Medicine. Retrieved from https://www.biospace.com/pixium-vision-announces-fda-approval-of-clinical-trial-expansion-of-prima-system-us-feasibility-study-to-stanford-medicine, 2022.

- Science Submits CE Mark Application for PRIMA Retinal Implant–A Critical Step Towards Making It Available To Patients. Retrieved from https://science.xyz/news/prima-ce-mark-application-submitted/, 2025.

- Sience Announces Positive Preliminary Results For Vision Restoration In Pivotal Clinical Trial. Retrieved from https://science.xyz/news/primavera-trial-preliminary-results/, 2024.

- Palanker D, Le Mer Y, et al. Photovoltaic Restoration of Central Vision in Atrophic Age-Related Macular Degeneration[J]. Ophthalmology, 2020, 127:1097-1104.

- Lemoine D, Simon E, Buc G, et al. In vitro reliability testing and in vivo lifespan estimation of wireless Pixium Vision PRIMA photovoltaic subretinal prostheses suggest prolonged durability and functionality in clinical practice[J]. Journal of neural engineering, 2020, 17(3):035005.

- Yang Rui, Zhao Peng, Wang Liyang, et al. Assessment of visual function in blind mice and monkeys with subretinally implanted nanowire arrays as artificial photoreceptors[J]. Nature Biomedical Engineering, 2024, 8(8):37.

- Wang Shuiyuan, Jiang Chengyong, Yu Yiye, et al. Tellurium nanowire retinal nanoprosthesis improves vision in models of blindness[J]. Science, 2025, 388(6751):eadu2987.

- Ayton L N, et al. First-in-human trial of a novel suprachoroidal retinal prosthesis[J]. PLoS One 9, 2014, 9(12):e115239.

- Fujikado T, et al. Testing of semichronically implanted retinal prosthesis by suprachoroidal-transretinal stimulation in patients with retinitis pigmentosa[J]. Investigative Ophthalmology & Visual Science, 2011, 52(7):4726-4733.

- Fujikado T, et al. One-Year Outcome of 49-Channel Suprachoroidal-Transretinal Stimulation Prosthesis in Patients With Advanced Retinitis Pigmentosa[J]. Investigative Ophthalmology & Visual Science, 2016, 57(14):6147-6157.

- Niketeghad S, Pouratian N, et al. Brain machine interfaces for vision restoration: the current state of cortical visual prosthetics[J]. Neurotherapeutics: the journal of the American Society for Experimental NeuroTherapeutics, 2018, 16:134–143.

- Barry M P, Armenta Salas M, Patel U, et al. Video-mode percepts are smaller than sums of single-electrode phosphenes with the Orion visual cortical prosthesis[J]. Investigative Ophthalmology & Visual Science, 2020, 61(6):3.

- Troyk P R. The intracortical visual prosthesis project[J]. Artificial Vision, 2017.

- Illinois Institute of Technology. First Successful Implantation of Revolutionary Wireless Visual Prosthesis Brain Implant. Retrieved from https://www.iit.edu/news/first-successful-implantation-revolutionary-wireless-visual-prosthesis-brain-implant, 2022.

- Lowery A J, Rosenfeld J V, Lewis P M, et al. Restoration of vision using wireless cortical implants: the monash vision group project[C]. IEEE Engineering in Medicine and Biology Society (EMBC), 2015.

- Lowery A J, Rosenfeld J V, Rosa M G P, et al. Monash vision group’s gennaris cortical implant for vision restoration[J]. Springer International Publishing, 2017, 215–225.

- Rosenfeld J V, Wong Y T, Yan E, et al. Tissue response to a chronically implantable wireless, intracortical visual prosthesis (Gennaris array)[J]. Journal of Neural Engineering, 2020, 17(4):046001.

- Gennaris. Retrieved from https://www.monash.edu/bioniceye/technology.

- Waisberg E. The potential to restore vision with Neuralink's "Blindsight" neural interface technology[J]. Pan-American Journal of Ophthalmology, 2024, 6.

- A new implant for blind people jacks directly into the brain. Retrieved from https://www.technologyreview.com/2020/02 /06/844908/a-new-implant-for-blind-people-jacks-directly-into-the-brain/2020-02-06.

- Beauchamp M S, Oswalt D, Sun P, et al. Dynamic Stimulation of Visual Cortex Produces Form Vision in Sighted and Blind Humans[J]. Cell, 2020, 181(4):774-783.

- Chen X, Wang F, Fernandez E, Roelfsema P R. Shape perception via a high-channel-count neuroprosthesis in monkey visual cortex[J]. Science, 2020, 370(6521):1191-1196.

- Fernandez E, Alfaro A, et al. Visual percepts evoked with an intracortical 96-channel microelectrode array inserted in human occipital cortex[J]. Journal of Clinical Investigation, 2021, 131(23):e151331.

- Sandoval-Pistorius S S, Hacker M L, Waters A C, et al. Advances in deep brain stimulation: From mechanisms to applications[J]. The Journal of Neuroscience, 2023, 43(45):7575-7586.

- Deuschl G, Schade-Brittinger C, Krack P, et al. A randomized trial of deep-brain stimulation for parkinson’s disease[J]. New England Journal of Medicine, 2006, 355(9):896-908.

- Wong J K, Hess C W, Almeida L, et al. Deep brain stimulation in essential tremor: Targets, technology, and a comprehensive review of clinical outcomes[J]. Expert Review of Neurotherapeutics, 2020, 20(4):319-331.

- Fan H, Zheng Z, Yin Z, et al. Deep brain stimulation treating dystonia: A systematic review of targets, body distributions and etiology classifications[J]. Frontiers in Human Neuroscience, 2021, 15:757579.

- Jia F, Shukla A W, Hu W, et al. Variable frequency deep brain stimulation of subthalamic nucleus to improve freezing of gait in parkinson’s disease[J]. National Science Review, 2024, 11(6):nwae187.

- Anderson D, Beecher G, Ba F. Deep brain stimulation in parkinson’s disease: New and emerging targets for refractory motor and nonmotor symptoms[J]. Parkinson’s Disease, 2017, 2017:1-13.

- Kühn A A, Volkmann J. Innovations in deep brain stimulation methodology[J]. Movement Disorders, 2017, 32(1):11-19.

- An Q, Yin Z, Ma R, et al. Adaptive deep brain stimulation for parkinson’s disease: Looking back at the past decade on motor outcomes[J]. Journal of Neurology, 2023, 270(3):1371-1387.

- Fisher R S, Boas W V E, Blume W, et al. Epileptic seizures and epilepsy: Definitions proposed by the international league against epilepsy (ILAE) and the international bureau for epilepsy (IBE)[J]. Epilepsia, 2005, 46(4):470-472.

- 中国抗癫痫协会, 中华医学会神经外科学分会神经生理学组, 中华医学会神经病学分会癫痫与脑电图学组等. 癫痫外科术前评估中国专家共识(2022版)[J]. 中华神经外科杂志, 2022, 38(10):973-979.

- Chen Z, Brodie M J, Liew D, et al. Treatment outcomes in patients with newly diagnosed epilepsy treated with established and new antiepileptic drugs: A 30-year longitudinal cohort study[J]. JAMA Neurology, 2018, 75(3):279-286.

- Jehi L, Jette N, Kwon C, et al. Timing of referral to evaluate for epilepsy surgery: Expert consensus recommendations from the surgical therapies commission of the international league against epilepsy[J]. Epilepsia, 2022, 63(10):2491-2506.

- Lamberink H J, Otte W M, Blümcke I, et al. Seizure outcome and use of antiepileptic drugs after epilepsy surgery according to histopathological diagnosis: A retrospective multicentre cohort study[J]. The Lancet Neurology, 2020, 19(9):748-757.

- Haneef Z, Skrehot H C. Neurostimulation in generalized epilepsy: A systematic review and meta‐analysis[J]. Epilepsia, 2023, 64(4):811-820.

- Touma L, Dansereau B, Chan A Y, et al. Neurostimulation in people with drug-resistant epilepsy: Systematic review and meta-analysis from the ILAE surgical therapies commission[J]. Epilepsia, 2022, 63(6):1314-1329.

- Toffa D H, Touma L, El Meskine T, et al. Learnings from 30 years of reported efficacy and safety of vagus nerve stimulation (VNS) for epilepsy treatment: A critical review[J]. Seizure, 2020, 83:104-123.

- Li M C H, Cook M J. Deep brain stimulation for drug‐resistant epilepsy[J]. Epilepsia, 2018, 59(2):273-290.

- Vetkas A, Fomenko A, Germann J, et al. Deep brain stimulation targets in epilepsy: Systematic review and meta‐analysis of anterior and centromedian thalamic nuclei and hippocampus[J]. Epilepsia, 2022, 63(3):513-524.

- Piper R J, Richardson R M, Worrell G, et al. Towards network-guided neuromodulation for epilepsy[J]. Brain, 2022, 145(10):3347-3362.

- Ngan Kee N, Foster E, Marquina C, et al. Systematic review of cost-effectiveness analysis for surgical and neurostimulation treatments for drug-resistant epilepsy in adults[J]. Neurology, 2023, 100(18).

- Nair D R, Laxer K D, Weber P B, et al. Nine-year prospective efficacy and safety of brain-responsive neurostimulation for focal epilepsy[J]. Neurology, 2020, 95(9).

- Liu J, Ning W, Zhang N, et al. Estimation of the global disease burden of depression and anxiety between 1990 and 2044: An analysis of the global burden of disease study 2019[J]. Healthcare, 2024, 12(17):1721.

- Lu J, Xu X, Huang Y, et al. Prevalence of depressive disorders and treatment in China: a cross-sectional epidemiological study[J]. The Lancet Psychiatry, 2021, 8(11):981-990.

- McIntyre R S, Alsuwaidan M, Baune B T, et al. Treatment-resistant depression: Definition, prevalence, detection, management, and investigational interventions[J]. World psychiatry: official journal of the World Psychiatric Association (WPA), 2023, 22(3):394-412.

- Hirschtritt M E, Bloch M H, Mathews C A. Obsessive-compulsive disorder: Advances in diagnosis and treatment[J]. JAMA, 2017, 317(13):1358.

- Huang Y, Wang Y, Wang H, et al. Prevalence of mental disorders in China: A cross-sectional epidemiological study[J]. The Lancet Psychiatry, 2019, 6(3):211-224.

- Alonso P, Segalàs C, Real E, et al. Suicide in patients treated for obsessive–compulsive disorder: A prospective follow-up study[J]. Journal of Affective Disorders, 2010, 124(3):300-308.

- Yang W, Tang Z, Wang X, et al. The cost of obsessive-compulsive disorder (OCD) in China: A multi-center cross-sectional survey based on hospitals[J]. General Psychiatry, 2021, 34(6): e100632.

- Bandelow B, Allgulander C, Baldwin D S, et al. World federation of societies of biological psychiatry (WFSBP) guidelines for treatment of anxiety, obsessive-compulsive and posttraumatic stress disorders – version 3. Part I: Anxiety disorders[J]. The World Journal of Biological Psychiatry, 2023, 24(2):79-117.

- 中华医学会精神医学分会《中国强迫症防治指南》编写组. 中国强迫症防治指南2016（精编版）[J]. 中华精神科杂志, 2016, 49(6):353-366.

- Howes O D, Thase M E, Pillinger T. Treatment resistance in psychiatry: State of the art and new directions[J]. Molecular Psychiatry, 2022, 27(1):58-72.

- Lozano A M, Lipsman N, Bergman H, et al. Deep brain stimulation: Current challenges and future directions[J]. Nature Reviews Neurology, 2019, 15(3):148-160.

- Gadot R, Najera R, Hirani S, et al. Efficacy of deep brain stimulation for treatment-resistant obsessive-compulsive disorder: Systematic review and meta-analysis[J]. Journal of Neurology, Neurosurgery & Psychiatry, 2022, 93(11):1166-1173.

- Kahn L, Sutton B, Winston H R, et al. Deep brain stimulation for obsessive-compulsive disorder: Real world experience post-FDA-humanitarian use device approval[J]. Frontiers in Psychiatry, 2021, 12:568932.

- US Food and Drug Administration. Deep brain stimulator for obsessive compulsive disorder[EB/OL]. Retrieved from https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfhde/hde.cfm?id=375533, 2025.

- Nuttin B, Wu H, Mayberg H, et al. Consensus on guidelines for stereotactic neurosurgery for psychiatric disorders[J]. Journal of Neurology, Neurosurgery & Psychiatry, 2014, 85(9):1003-1008.

- Hamani C, Pilitsis J, Rughani A I, et al. Deep brain stimulation for obsessive-compulsive disorder: Systematic review and evidence-based guideline sponsored by the American society for stereotactic and functional neurosurgery and the congress of neurological surgeons (CNS) and endorsed by the CNS and american association of neurological surgeons[J]. Neurosurgery, 2014, 75(4):327-333.

- 中华医学会神经外科分会功能神经外科学组, 中国医师协会精神科医师分会精神疾病外科治疗联盟, 中国医师协会神经调控专业委员会精神疾病学组. 难治性强迫症脑深部电刺激疗法中国专家共识[J]. 临床神经外科杂志, 2024, 21(5):481-487,494.

- Meyer G M, Hollunder B, Li N, et al. Deep brain stimulation for obsessive-compulsive disorder: Optimal stimulation sites[J]. Biological Psychiatry, 2024, 96(2):101-113.

- Edlow B L, Claassen J, Schiff N D, et al. Recovery from disorders of consciousness: mechanisms, prognosis and emerging therapies[J]. Nature Reviews Neurology, 2021, 17(3):135-156.

- Lulé D, Noirhomme Q, Kleih S C, et al. Probing command following in patients with disorders of consciousness using a brain–computer interface[J]. Clinical Neurophysiology, 2013, 124(1):101-106.

- Chatelle C, Lesenfants D, Guller Y, et al. Brain-computer interface for assessing consciousness in severely brain-injured patients[J]. Clinical neurophysiology in disorders of consciousness: Brain function monitoring in the ICU and beyond, 2015:133-148.

- Liuzzi P, Grippo A, Campagnini S, et al. Merging clinical and EEG biomarkers in an elastic-net regression for disorder of consciousness prognosis prediction[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2022, 30:1504-1513.

- Schiff N D, Diringer M, Diserens K, et al. Brain–computer interfaces for communication in patients with disorders of consciousness: A gap analysis and scientific roadmap[J]. Neurocritical Care, 2024, 41(1):129-145.

- He Q, Yang Y, Ge P, et al. The brain nebula: minimally invasive brain–computer interface by endovascular neural recording and stimulation[J]. Journal of NeuroInterventional Surgery, 2024, 16(12):1237-1243.

- Curley W H, Forgacs P B, Voss H U, et al. Characterization of EEG signals revealing covert cognition in the injured brain[J]. Brain, 2018, 141(5):1404-1421.

- Claassen J, Kondziella D, Alkhachroum A, et al. Cognitive motor dissociation: gap analysis and future directions[J]. Neurocritical care, 2024, 40(1):81-98.

- Wu M, Bao W-X, Zhang J, et al. Effect of acoustic stimuli in patients with disorders of consciousness: a quantitative electroencephalography study[J]. Neural regeneration research, 2018, 13(11):1900-1906.

- Claassen J, Doyle K, Matory A, et al. Detection of brain activation in unresponsive patients with acute brain injury[J]. New England Journal of Medicine, 2019, 380(26):2497-2505.

- Gui P, Jiang Y, Zang D, et al. Assessing the depth of language processing in patients with disorders of consciousness[J]. Nature neuroscience, 2020, 23(6):761-770.

- Spataro R, Heilinger A, Allison B, et al. Preserved somatosensory discrimination predicts consciousness recovery in unresponsive wakefulness syndrome[J]. Clinical Neurophysiology, 2018, 129(6):1130-1136.

- Pan J, Xie Q, Qin P, et al. Prognosis for patients with cognitive motor dissociation identified by brain-computer interface[J]. Brain, 2020, 143(4):1177-1189.

- Xu R, Spataro R, Allison B Z, et al. Brain–computer interfaces in acute and subacute disorders of consciousness[J]. Journal of Clinical Neurophysiology, 2022, 39(1):32-39.

- Vatrano M, Nemirovsky I E, Tonin P, et al. Assessing consciousness through neurofeedback and neuromodulation: possibilities and challenges[J]. Life, 2023, 13(8):1675.

- Stender J, Gosseries O, Bruno M-A, et al. Diagnostic precision of PET imaging and functional MRI in disorders of consciousness: a clinical validation study[J], Lancet, 2014, 384(9942):514-522.

- Kondziella D, Bender A, Diserens K, et al. European Academy of Neurology guideline on the diagnosis of coma and other disorders of consciousness[J]. European journal of neurology, 2020, 27(5):741-756.

- Giacino J T, Katz D I, Schiff N D, et al. Practice guideline update recommendations summary: disorders of consciousness: report of the Guideline Development, Dissemination, and Implementation Subcommittee of the American Academy of Neurology; the American Congress of Rehabilitation Medicine; and the National Institute on Disability, Independent Living, and Rehabilitation Research[J]. Neurology, 2018, 91(10):450-460.

- Bahammam A S. Artificial Intelligence in Sleep Medicine: The Dawn of a New Era[J]. Nature and Science of Sleep, 2024:445-450.

- Supratak A, Dong H, Wu C, et al. DeepSleepNet: a Model for Automatic Sleep Stage Scoring based on Raw Single-Channel EEG[J]. IEEE Transactions on Neural Systems & Rehabilitation Engineering A Publication of the IEEE Engineering in Medicine & Biology Society, 2017, (99):1-1.

- Phan H, Chén O Y, Tran M C, et al. XSleepNet: Multi-view sequential model for automatic sleep staging[J]. IEEE Transactions on Pattern Analysis and Machine Intelligence, 2021, 44(9):5903-5915.

- Zhu T, Luo W, Yu F. Convolution-and attention-based neural network for automated sleep stage classification[J]. International Journal of Environmental Research and Public Health, 2020, 17(11):4152.

- Jia Z, Lin Y, Wang J, et al. GraphSleepNet: Adaptive spatial-temporal graph convolutional networks for sleep stage classification[C]. International Joint Conference on Artificial Intelligence, 2020:1324-1330.

- You Y, Zhong X, Liu G, et al. Automatic sleep stage classification: A light and efficient deep neural network model based on time, frequency and fractional Fourier transform domain features[J]. Artificial Intelligence in Medicine, 2022, 127:102279.

- Lachner-Piza D, Epitashvili N, Schulze-Bonhage A, et al. A single channel sleep-spindle detector based on multivariate classification of EEG epochs: MUSSDET[J]. Journal of neuroscience methods, 2018, 297:31-43.

- Chen P, Chen D, Zhang L, et al. Automated sleep spindle detection with mixed EEG features[J]. Biomedical Signal Processing and Control, 2021, 70:103026.

- 李小俚, 王枫, 黄朝阳等. 深度学习的睡眠脑电特征波检测[J]. 北京师范大学学报 (自然科学版), 2021, 57(6):860-867.

- Jiang D, Ma Y, Wang Y. A robust two-stage sleep spindle detection approach using single-channel EEG[J]. Journal of Neural Engineering, 2021, 18(2):026026.

- Chen C, Meng J, Belkacem A N, et al. Hierarchical fusion detection algorithm for sleep spindle detection[J]. Frontiers in Neuroscience, 2023, 17:1105696.

- Lanza G, Cantone M, Lanuzza B, et al. Distinctive patterns of cortical excitability to transcranial magnetic stimulation in obstructive sleep apnea syndrome, restless legs syndrome, insomnia, and sleep deprivation[J]. Sleep medicine reviews, 2015, 19:39-50.

- Krause M R, Vieira P G, Csorba B A, et al. Transcranial alternating current stimulation entrains single-neuron activity in the primate brain[J]. Proceedings of the National Academy of Sciences, 2019, 116(12):5747-5755.

- Xie J, Wang L, Xiao C, et al. Low frequency transcranial alternating current stimulation accelerates sleep onset process[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2021, 29:2540-2549.

- 罗帏, 周仁来. 经颅电刺激技术在睡眠调节中的作用和应用展望[J]. 航天医学与医学工程, 2020, 33(2):7..

- Li R, Yang D, Fang F, et al. Concurrent fNIRS and EEG for Brain Function Investigation: A Systematic, Methodology-Focused Review[J]. Sensors (Basel), 2022, 22(15).

- Girouard H, Iadecola C. Neurovascular coupling in the normal brain and in hypertension, stroke, and Alzheimer disease[J]. Journal of Applied Physiology, 2006, 100(1):328-35.

- Li C, Tang T, Pan Y, et al. An efficient graph learning system for emotion recognition inspired by the cognitive prior graph of EEG brain network[J], 2024, 36.

- Liu F, Yang P, Shu Y, et al. Emotion recognition from few-channel eeg signals by integrating deep feature aggregation and transfer learning[J], 2023, 15(3):1315-1330.

- Xu L, Ma Z, Meng J, et al. Improving transfer performance of deep learning with adaptive batch normalization for brain-computer interfaces[C]. Annual International Conference of the IEEE Engineering in Medicine & Biology Society (EMBC), 2021:5800-5803.

- Cai H, Pan J, Xiao Q, et al. Decoding Musical Neural Activity in Patients With Disorders of Consciousness Through Self-Supervised Contrastive Domain Generalization[J]. IEEE Transactions on Affective Computing, 2024.

- Yang C, Westover M, et al. Biot: Biosignal transformer for cross-data learning in the wild[C]. Conference and Workshop on Neural Information Processing Systems, 2023, 36:78240-78260.

- Lee S, Kim H, Park M J, et al. Current advances in wearable devices and their sensors in patients with depression[J]. Frontiers in psychiatry, 2021, 12:672347.

- Qi Z, Liu H, Jin F, et al. A wearable repetitive transcranial magnetic stimulation device[J]. Nature Communications, 2025, 16(1):2731.

- Lin C-T, Lin B-S, Lin F-C, et al. Brain computer interface-based smart living environmental auto-adjustment control system in UPnP home networking[J]. IEEE Systems Journal, 2014, 8(2):363-370.

- Zhu J, Bao X, Huang Q, et al. Motion Sickness Alleviation Based on a Mindfulness Brain-Computer Interface[C]. International Conference on Advanced Computational Intelligence (ICACI), 2024:1-6.

- Li G, Zhang L, Li C, et al. Investigating the Influence of Driving on Brain Connectivity Networks and Emotion Processing Mechanism Based on EEG Signals[J]. IEEE Sensors Journal, 2024, 21:24.

- Hofmann S M, Klotzsche F, Mariola A, et al. Decoding subjective emotional arousal from EEG during an immersive virtual reality experience[J]. eLife Sciences Publications Limited, 2021, 10:e64812.

- Jung T-P, Makeig S, Stensmo M, et al. Estimating alertness from the EEG power spectrum[J]. IEEE transactions on biomedical engineering, 1997, 44(1):60-69.

- Li W, Ming D, Xu R, et al. Research on visual attention classification based on EEG entropy parameters[C]. World Congress on Medical Physics and Biomedical Engineering May 26-31, 2012, Beijing, China, 2013:1553-1556.

- Wang Y-K, Jung T-P, Lin C-T. EEG-based attention tracking during distracted driving[J]. IEEE transactions on neural systems and rehabilitation engineering, 2015, 23(6):1085-1094.

- Hu B, Li X, Sun S, et al. Attention recognition in EEG-based affective learning research using CFS+ KNN algorithm[J]. IEEE/ACM transactions on computational biology and bioinformatics, 2016, 15(1):38-45.

- Chen D, Huang H, Guan Z, et al. An Intersubject Brain-Computer Interface Based on Domain-Adversarial Training of Convolutional Neural Network[J]. IEEE Transactions on Biomedical Engineering, 2024, 71(10):2956-2967.

- Värbu K, Muhammad N, Muhammad Y. Past, present, and future of EEG-based BCI applications[J]. Sensors, 2022, 22(9):3331.

- Lim C G, Soh C P, Lim S S Y, et al. Home-based brain–computer interface attention training program for attention deficit hyperactivity disorder: A feasibility trial[J]. Child and Adolescent Psychiatry and Mental Health, 2023, 17(1):15.

- Huang H, Chen J, Xiao J, et al. Real-Time Attention Regulation and Cognitive Monitoring Using a Wearable EEG-Based BCI[J]. IEEE Transactions on Biomedical Engineering, 2025, 72(2):716-724.

- Jiang J, Yan Z, Sheng C, et al. A Novel Detection Tool for Mild Cognitive Impairment Patients Based on Eye Movement and Electroencephalogram[J]. Journal of Alzheimer's Disease, 2019, 72(2):389-399.

- Bakhshayesh A R, Hänsch S, Wyschkon A, et al. Neurofeedback in ADHD: a single-blind randomized controlled trial[J]. European Child & Adolescent Psychiatry, 2011, 20(9): 481-491.

- Zang Y-F, Lim C G, Lee T S, et al. A Brain-Computer Interface Based Attention Training Program for Treating Attention Deficit Hyperactivity Disorder[J]. PLoS ONE, 2012, 7(10).

- Pliszka S, Issues A W G O Q. Practice parameter for the assessment and treatment of children and adolescents with attention-deficit/hyperactivity disorder[J]. Journal of the American Academy of Child & Adolescent Psychiatry, 2007, 46(7):894-921.

- Li Z J, Zhang L B, Chen Y X, et al. Advancements and challenges in neuromodulation technology: interdisciplinary opportunities and collaborative endeavors[J]. Science Bulletin, 2023, 68(18):1978-1982.

- Topka M, Schneider M, Zrenner C, et al. Motor cortex excitability is reduced during freezing of upper limb movement in Parkinson’s disease[J]. npj Parkinson's Disease, 2022, 8(1):161.

- Zeng K, Darmani G, Fomenko A, et al. Induction of human motor cortex plasticity by theta burst transcranial ultrasound stimulation[J]. Annals of neurology, 2022, 91(2):238-252.

- Attal N, Poindessous-Jazat F, De Chauvigny E, et al. Repetitive transcranial magnetic stimulation for neuropathic pain: a randomized multicentre sham-controlled trial[J]. Brain, 2021, 144(11):3328-3339.

- Koch G, Casula E P, Bonnì S, et al. Precuneus magnetic stimulation for Alzheimer’s disease: a randomized, sham-controlled trial[J]. Brain, 2022, 145(11):3776-3786.

- Noda Y, Osawa R, Takeda Y, et al. Left prefrontal intermittent theta-burst stimulation therapy for major depressive disorder: A real-world, multisite observational study in Japan[J]. Journal of Affective Disorders, 2025, 375:316-323.

- Wu X, Ji G J, Geng Z, et al. Accelerated intermittent theta-burst stimulation broadly ameliorates symptoms and cognition in Alzheimer's disease: A randomized controlled trial[J]. Brain Stimulation, 2022, 15(1):35-45.

- Zangen A, Zibman S, Tendler A, et al. Pursuing personalized medicine for depression by targeting the lateral or medial prefrontal cortex with Deep TMS[J]. JCI insight, 2023, 8(4):e165271.

- Smárason O, Boedeker P J, Guzick A G, et al. Depressive symptoms during deep transcranial magnetic stimulation or sham treatment for obsessive-compulsive disorder[J]. Journal of Affective Disorders, 2024, 344:466-472.

- Anli L, Mihály V, Greg K, et al. Immediate neurophysiological effects of transcranial electrical stimulation[J]. Nature Communications, 2018, 9(1):5092.

- Rafael P, A M N, C C R. Studying and modifying brain function with non-invasive brain stimulation[J]. Nature neuroscience, 2018, 21(2):174-187.

- Loo C K, Alonzo A, Martin D, et al. Transcranial direct current stimulation for depression: 3-week, randomised, sham-controlled trial[J]. The British Journal of Psychiatry, 2012, 200(1):52-59.

- Brunelin J, Mondino M, Gassab L, et al. Examining transcranial direct-current stimulation (tDCS) as a treatment for hallucinations in schizophrenia[J]. American Journal of Psychiatry, 2012, 169(7):719–24.

- Valiengo L, Goerigk S, Gordon PC, et al. Efficacy and safety of transcranial direct current stimulation for treating negative symptoms in schizophrenia: a randomized clinical trial[J]. JAMA Psychiatry, 2020, 77(2):121–9.

- Hyde J, Carr H, Kelley N, et al. Efficacy of neurostimulation across mental disorders: systematic review and meta-analysis of 208 randomized controlled trials[J]. Molecular Psychiatry, 2022, 27(6): 2709-2719.

- Alberto B, Valentina D, Valentina C, et al. Cerebello-spinal tDCS in ataxia: A randomized, double-blind, sham-controlled, crossover trial[J]. Neurology, 2018, 91(12):e1090-e1101.

- Van’t Wout-Frank M, Arulpragasam A R, Faucher C, et al. Virtual reality and transcranial direct current stimulation for posttraumatic stress disorder: a randomized clinical trial[J]. JAMA psychiatry, 2024, 81(5):437-446.

- H F K, Katharina D, C M M, et al. Integrating electric field modeling and neuroimaging to explain inter-individual variability of tACS effects[J]. Nature Communications, 2019, 10(1):5427.

- Alekseichuk I, Turi Z, Lara D A G, et al. Spatial Working Memory in Humans Depends on Theta and High Gamma Synchronization in the Prefrontal Cortex[J]. Current Biology, 2016, 26(12):1513-1521.

- Alberto B, Valentina C, Sofia M C, et al. Exposure to gamma tACS in Alzheimer’s disease: A randomized, double-blind, sham-controlled, crossover, pilot study[J]. Brain Stimulation, 2021, 14(3):531-540.

- Dongsheng Z, Ang L, Xingxing L, et al. Effects of 40 Hz transcranial alternating current stimulation (tACS) on cognitive functions of patients with Alzheimer's disease: a randomised, double-blind, sham-controlled clinical trial[J]. Journal of Neurology, Neurosurgery, and Psychiatry, 2021.

- Brittain J, Probert-Smith P, Aziz Z T, et al. Tremor Suppression by Rhythmic Transcranial Current Stimulation[J]. Current Biology, 2013, 23(5):436-440.

- Shrey G, A J N, Vighnesh V, et al. High-frequency neuromodulation improves obsessive-compulsive behavior[J]. Nature Medicine, 2021, 27(2):232-238.

- Van der Groen O, Potok W, Wenderoth N, et al. Using noise for the better: The effects of transcranial random noise stimulation on the brain and behavior[J]. Neuroscience & Biobehavioral Reviews, 2022, 138:104702.

- Van Der Groen O, Wenderoth N. Transcranial random noise stimulation of visual cortex: stochastic resonance enhances central mechanisms of perception[J]. Journal of Neuroscience, 2016, 36(19):5289-5298.

- Murphy O W, Hoy K E, Wong D, et al. Transcranial random noise stimulation is more effective than transcranial direct current stimulation for enhancing working memory in healthy individuals: Behavioural and electrophysiological evidence[J]. Brain stimulation, 2020, 13(5):1370-1380.

- Grossman N, Bono D, Dedic N, et al. Noninvasive Deep Brain Stimulation via Temporally Interfering Electric Fields[J]. Cell, 2017, 169(6):1029-1041.e16.

- Wessel M J, Beanato E, Popa T, et al. Noninvasive theta-burst stimulation of the human striatum enhances striatal activity and motor skill learning[J]. Nature neuroscience, 2023, 26(11):2005-2016.

- Violante I R, Alania K, Cassarà A M, et al. Non-invasive temporal interference electrical stimulation of the human hippocampus[J]. Nature neuroscience, 2023, 26(11):1994-2004.

- Yang C H, Xu Y X, Feng X F, et al. Transcranial Temporal Interference Stimulation of the Right Globus Pallidus in Parkinson's Disease[J]. Movement Disorders, 2024, 40(6).

- Beanato E, Moon H J, Windel F, et al. Noninvasive modulation of the hippocampal-entorhinal complex during spatial navigation in humans[J]. Science Advances, 2024, 10(44).

- Meng Y, Hynynen K, Lipsman N. Applications of focused ultrasound in the brain: from thermoablation to drug delivery[J]. Nature Reviews Neurology, 2021, 17(1):7-22.

- Folloni D, Verhagen L, Mars R B, et al. Manipulation of subcortical and deep cortical activity in the primate brain using transcranial focused ultrasound stimulation[J]. Neuron, 2019, 101(6):1109-1116.e5.

- Lipsman N, Meng Y, Bethune A J, et al. Blood–brain barrier opening in Alzheimer’s disease using MR-guided focused ultrasound[J]. Nature communications, 2018, 9(1):2336.

- Rezai A R, Ranjan M, D’Haese P F, et al. Noninvasive hippocampal blood− brain barrier opening in Alzheimer’s disease with focused ultrasound[J]. Proceedings of the National Academy of Sciences, 2020, 117(17):9180-9182.

- Abrahao A, Meng Y, Llinas M, et al. First-in-human trial of blood–brain barrier opening in amyotrophic lateral sclerosis using MR-guided focused ultrasound[J]. Nature communications, 2019, 10(1):4373.

- Legon W, Strohman A. Low-intensity focused ultrasound for human neuromodulation[J]. Nature Reviews Methods Primers, 2024, 4(1):91.

- Samuel N, Ding M Y R, Sarica C, et al. Accelerated transcranial ultrasound neuromodulation in Parkinson's disease: a pilot study[J]. Movement Disorders, 2023, 38(12):2209-2216.

- Bancel T, Béranger B, Daniel M, et al. Sustained reduction of essential tremor with low-power non-thermal transcranial focused ultrasound stimulations in humans[J]. Brain Stimulation, 2024, 17(3):636-647.

- Ziebell P, Rodrigues J, Forster A, et al. Inhibition of midfrontal theta with transcranial ultrasound explains greater approach versus withdrawal behavior in humans[J]. Brain Stimulation, 2023, 16(5):1278-1288.

- Fan J M, Woodworth K, Murphy K R, et al. Thalamic transcranial ultrasound stimulation in treatment resistant depression[J]. Brain stimulation, 2024, 17(5):1001.

- Beisteiner R, Matt E, Fan C, et al. Transcranial pulse stimulation with ultrasound in Alzheimer's disease—a new navigated focal brain therapy[J]. Advanced Science, 2020, 7(3):1902583.

- Lee C C, Chou C C, Hsiao F J, et al. Pilot study of focused ultrasound for drug‐resistant epilepsy[J]. Epilepsia, 2022, 63(1):162-175.

- Stern J M, Spivak N M, Becerra S A, et al. Safety of focused ultrasound neuromodulation in humans with temporal lobe epilepsy[J]. Brain stimulation, 2021, 14(4):1022-1031.

- Riis T S, Feldman D A, Losser A J, et al. Noninvasive targeted modulation of pain circuits with focused ultrasonic waves[J]. Pain, 2024, 165(12):2829-2839.

- Mahoney J J, Thompson-Lake D G Y, Ranjan M, et al. Low-intensity focused ultrasound targeting the bilateral nucleus accumbens as a potential treatment for substance use disorder: a first-in-human report[J]. Biological Psychiatry, 2023, 94(11):e41-e43.

- Sarica C, Nankoo J F, Fomenko A, et al. Human studies of transcranial ultrasound neuromodulation: a systematic review of effectiveness and safety[J]. Brain stimulation, 2022, 15(3):737-746.

- Wang L, Mao L, Huang Z, et al. Photobiomodulation: shining a light on depression[J]. Theranostics, 2025, 15(2):362.

- Hamblin M R. Shining light on the head: photobiomodulation for brain disorders[J]. BBA clinical, 2016, 6:113-124.

- Lee T, Ding Z, Chan A S. Can transcranial photobiomodulation improve cognitive function? A systematic review of human studies[J]. Ageing research reviews, 2023, 83:101786.

- Iaccarino H F, Singer A C, Martorell A J, et al. Gamma frequency entrainment attenuates amyloid load and modifies microglia[J]. Nature, 2016, 540(7632):230-235.

- Zhao C, Li D, Kong Y, et al. Transcranial photobiomodulation enhances visual working memory capacity in humans[J]. Science advances, 2022, 8(48):eabq3211.

- X Qu, Li L, Zhou X, et al. Repeated transcranial photobiomodulation improves working memory of healthy older adults: behavioral outcomes of poststimulation including a three-week follow-up[J]. Neurophotonics, 2022, 9(3):035005-035005.

- Herkes G, McGee C, Liebert A, et al. A novel transcranial photobiomodulation device to address motor signs of Parkinson's disease: a parallel randomised feasibility study[J]. EClinicalMedicine, 2023, 66.

- Wang L, Mao L, Huang Z, et al. Photobiomodulation: shining a light on depression[J]. Theranostics, 2025, 15(2):362.

- GALÁN F, NUTTIN M, LEW E, et al. A brain-actuated wheelchair: Asynchronous and non-invasive brain–computer interfaces for continuous control of robots[J]. Clinical Neurophysiology, 2008, 119(9):2159-2169.

- Barsotti M, Leonardis D, Loconsole C, et al. A full upper limb robotic exoskeleton for reaching and grasping rehabilitation triggered by MI-BCI[C]. Proceedings of the IEEE/RAS-EMBS International Conference on Rehabilitation Robotics(ICORR), 2015, Singapore.

- BIASIUCCI A, LEEB R, ITURRATE I, et al. Brain-actuated functional electrical stimulation elicits lasting arm motor recovery after stroke[J]. Nature Communications, 2018, 9(1):2421.

- Ramos-Murguialday A, et al. Brain-Machine Interface in Chronic Stroke: Randomized Trial Long-Term Follow-up[J]. Neurorehabilitation and Neural Repair, 2019, 33(3):188-198.

- Keser Z, Buchl S C, Seven N A, et al. Electroencephalogram (EEG) with or without transcranial magnetic stimulation (TMS) as biomarkers for post-stroke recovery: a narrative review[J]. Frontiers in neurology, 2022, 13:827866.

- Lyukmanov R K, Isaev M R, Mokienko O A, et al. Brain-Computer Interface Using Functional Near-Infrared Spectroscopy for Post-Stroke Motor Rehabilitation: Case Series[J]. Annals of clinical and experimental neurology, 2023, 17(4):82-88.

- Brunner I, Lundquist C B, Pedersen A R, et al. Brain computer interface training with motor imagery and functional electrical stimulation for patients with severe upper limb paresis after stroke: A randomized controlled pilot trial[J]. Journal of NeuroEngineering and Rehabilitation, 2024, 21(1):10.

- Kim T, Wang Y, Cho H, et al. NOIR 2.0: Neural Signal Operated Intelligent Robots for Everyday Activities[C]. CoRL Workshop CoRoboLearn: Advancing Learning for Human-Centered Collaborative Robots, 2024.

- 陈树耿，束小康，贾杰. 基于闭环脑机接口的脑卒中患者的手功能康复研究[J]. 中国康复医学杂志，2016,31(11):1189-1194.

- 吴琼，任诗媛，乐赞等. 脑机接口综合康复训练对亚急性期脑卒中疗效的静息态功能磁共振研究[J]. 中国康复理论与实践，2020, 26(1):7784.

- Guo N, Wang X, Duanmu D, et al. SSVEP-based brain computer interface controlled soft robotic glove for post-stroke hand function rehabilitation[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2022, 30:1737-1744.

- 李叶. BrainRobotics智能仿生手[J]. 设计, 2022, 35(12):36-40.

- XU X, FAN X, DONG J, et al. Event-related EEG desynchronization reveals enhanced motor imagery from the third person perspective by manipulating sense of body ownership with virtual reality for stroke patients[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2024, 32:1055-1067.

- Z WANG, Y LIU, S HUANG, et al. Enhancing ERD activation and functional connectivity via the sixth-finger motor imagery in stroke patients[J]. IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2024, 32:3902-3912.

- Z WANG, Y LIU, S HUANG, et al. EEG characteristic comparison of motor imagery between supernumerary and inherent limb: Sixth-finger MI enhances the ERD pattern and classification performance[J]. IEEE Journal of Biomedical and Health Informatics, 2024, 28(12):7078-7089.

- Liu Y, Gui Z, Yan D, et al. Lower limb motor imagery EEG dataset based on the multi-paradigm and longitudinal-training of stroke patients[J]. Scientific Data, 2025, 12(1):314.